## Sheet1
Date | Site | SRP | SSH | sonicshow | irgashow | diskuse% | filesize | files | Notes
02 Jan 2023 | CH-FRU | LH | 0 | 0 | 0 | 35 | 0 | 0
30 Dec 2022 | CH-AWS | PM | 0 | 0 | 0 | 66 | 0 | 0
30 Dec 2022 | CH-CHA | PM | 0 | 0 | 1 | 41 | 0 | 0 | "Window 93.8% RAIN or NEEDS CLEANING", probably due to dew
30 Dec 2022 | CH-DAV | PM | 0 | 0 | 0 | 51 | 0 | 0
30 Dec 2022 | CH-DAS | PM | 0 | 0 | 0 | 75 | 0 | 0
30 Dec 2022 | CH-FRU | LH | 0 | 0 | 0 | 35 | 0 | 0
29 Dec 2022 | CH-CHA | YW | 0 | 0 | 0 | 39 | 0 | 0
20 Dec 2022 | CH-OE2 | FT | 0 | 0 | 0 | 13 | 0 | 0 | LGR cell pressure is decreasing too fast
20 Dec 2022 | CH-DAV | IF | 0 | 0 | 0 | 49 | 0 | 0
20 Dec 2022 | CH-LAS | LS | 0 | 0 | 0 | 60 | 1 | 0 | The 14th of December the file size is 19 from 9:00 until 14:00. Then the size increases again
20 Dec 2022 | CH-CHA | YW | 0 | 0 | 0 | 40 | 0 | 0
19 Dec 2022 | CH-CHA | YW | 0 | 0 | 0 | 78 | 0 | 0
19 Dec 2022 | CH-FRU | LH | 0 | 0 | 0 | 35 | 0 | 0
16 Dec 2022 | CH-AWS | GA | 0 | 0 | 0 | 64 | 0 | 0
13 Dec 2022 | CH-DAV | IF | 0 | 0 | 0 | 49 | 0 | 0 | Smaller file on 2022120913.X00 (29 KB) might be due to snowfall
13 Dec 2022 | CH-LAS | LS | 0 | 0 | 0 | 40 | 0 | 0
12 Dec 2022 | CH-CHA | YW | 0 | 0 | 0 | 77 | 0 | 0
12 Dec 2022 | CH-FRU | LH | 0 | 0 | 1 | 35 | 0 | 0 | at the moment likely ice on the sensor --> Window 93.8% RAIN or NEEDS CLEANING; on 4 Dec data gap of several hours likely because of the same issue
12 Dec 2022 | CH-OE2 | FT | 0 | 0 | 0 | 13 | 0 | 0
11 Dec 2022 | CH-AWS | GA | 0 | 0 | 0 | 63 | 0 | 0
09 Dec 2022 | CH-DAS | PM | 0 | 0 | 1 | 72 | 0 | 0 | elevated window dirtiness and unrealistic concentration data due to snow-fall
08 Dec 2022 | CH-OE2 | FT | 0 | 0 | 0 | 13 | 0 | 0
06 Dec 2022 | CH-DAS | LK | 0 | 0 | 0 | 71 | 0 | 0
06 Dec 2022 | CH-CHA | YW | 0 | 0 | 0 | 76 | 0 | 0
06 Dec 2022 | CH-CHA | PM | 0 | 0 | 0 | 88 | 0 | 0
06 Dec 2022 | CH-LAE | AS | 0 | 1 | 1 | 7 | 0 | 0 | Old IRGA flow flow-rate problem. Also AGC now 93.4%
06 Dec 2022 | CH-AWS | PM | 0 | 0 | 0 | 62 | 0 | 0
06 Dec 2022 | CH-LAS | LS | 0 | 0 | 0 | 59 | 0 | 0
05 Dec 2022 | CH-DAV | IF | 0 | 0 | 0 | 48 | 0 | 0
02 Dec 2022 | CH-LAS | TB | 0 | 0 | 0 | 58 | 0 | 0
02 Dec 2022 | CH-LAE | TB | 0 | 0 | 1 | 6 | 0 | 0
02 Dec 2022 | CH-AWS | GA | 0 | 0 | 0 | 61 | 0 | 0
02 Dec 2022 | CH-FRU | TB | 0 | 0 | 0 | 35 | 0 | 0
01 Dec 2022 | CH-DAV | IF | 0 | 0 | 0 | 47 | 0 | 1 | One slightly smaller file: 2022112907.X00 Could be related to station visit on the same day.
30 Nov 2022 | CH-OE2 | FT | 0 | 0 | 0 | 12 | 0 | 0
29 Nov 2022 | CH-DAS | LK | 0 | 0 | 0 | 70 | 0 | 0
28 Nov 2022 | CH-AWS | PM | 0 | 0 | 0 | 61 | 0 | 0
28 Nov 2022 | CH-CHA | PM | 0 | 0 | 1 | 79 | 0 | 0 | IRGA with negative H2O density, window dirtyness at 62.5%. Probably water droplet on the lens
28 Nov 2022 | CH-FRU | LH | 0 | 0 | 0 | 35 | 0 | 0
28 Nov 2022 | CH-CHA | YW | 0 | 0 | 0 | 75 | 0 | 0 | H2O back to positive value again.
25 Nov 2022 | CH-AWS | GA | 0 | 0 | 0 | 60 | 0 | 0
25 Nov 2022 | CH-LAS | TB | 0 | 0 | 0 | 57 | 0 | 0
25 Nov 2022 | CH-LAE | TB | 0 | 0 | 1 | 6 | 0 | 0
25 Nov 2022 | CH-FRU | TB | 0 | 0 | 0 | 35 | 0 | 0
24 Nov 2022 | CH-LAS | LS | 0 | 0 | 0 | 57 | 0 | 0
22 Nov 2022 | CH-OE2 | FT | 0 | 0 | 0 | 12 | 0 | 0
22 Nov 2022 | CH-CHA | YW | 0 | 0 | 0 | 74 | 0 | 0
22 Nov 2022 | CH-AWS | GA | 0 | 0 | 0 | 60 | 0 | 0
21 Nov 2022 | CH-DAS | LK | 0 | 0 | 0 | 69 | 0 | 0
21 Nov 2022 | CH-FRU | SO | 0 | 0 | 0 | 34 | 0 | 0 | Sonic temp shows about 12°C and is higher than irga temp (ca. 5°C), station temp (ca. 4°C) and Sattel, SZ temp (ca. 4°C).
21 Nov 2022 | CH-DAV | IF | 0 | 0 | 0 | 46 | 0 | 0
17 Nov 2022 | CH-FRU | SO | 0 | 0 | 0 | 34 | 0 | 0
15 Nov 2022 | CH-DAS | LK | 0 | 0 | 0 | 68 | 0 | 0
15 Nov 2022 | CH-LAS | LS | 0 | 0 | 0 | 56 | 0 | 0
15 Nov 2022 | CH-OE2 | FT | 0 | 0 | 0 | 12 | 0 | 0 | irga signal back to 100.1% after replacement of devicecell pressure of LGR is dropping --> need to change filters
14 Nov 2022 | CH-CHA | YW | 0 | 0 | 1 | 73 | 0 | 0 | irga window dirtiness high probably due to rain
14 Nov 2022 | CH-DAV | IF | 0 | 0 | 0 | 45 | 0 | 0
14 Nov 2022 | CH-FRU | LH | 0 | 0 | 0 | 34 | 0 | 0
11 Nov 2022 | CH-AWS | GA | 0 | 0 | 0 | 57 | 0 | 0
10 Nov 2022 | CH-LAE | AS | 0 | 1 | 1 | 6 | 0 | 0 | The same flow rate problem with IRGA (please see previous entry). The new IRGA will be installed soon and that should solve the problem.
08 Nov 2022 | CH-CHA | PM | 0 | 0 | 1 | 77 | 0 | 0 | IRGA window dirtyness high probably due to dew and fog
08 Nov 2022 | CH-DAS | PM | 0 | 0 | 0 | 67 | 0 | 0
08 Nov 2022 | CH-OE2 | FT | 0 | 0 | 1 | 12 | 0 | 0 | irga signal dropped from 100.1 to 66.7: need to understand if it is due to dew
08 Nov 2022 | CH-LAS | LS | 0 | 0 | 0 | 55 | 0 | 0
07 Nov 2022 | CH-CHA | YW | 0 | 0 | 0 | 72 | 0 | 0
07 Nov 2022 | CH-AWS | PM | 0 | 0 | 0 | 57 | 0 | 0
07 Nov 2022 | CH-FRU | LH | 0 | 0 | 0 | 34 | 0 | 0
07 Nov 2022 | CH-DAV | IF | 0 | 0 | 0 | 44 | 0 | 0
05 Nov 2022 | CH-AWS | GA | 0 | 0 | 0 | 57 | 0 | 0
04 Nov 2022 | CH-OE2 | TB | 0 | 0 | 0 | 12 | 0 | 0
04 Nov 2022 | CH-LAE | TB | 0 | 0 | 1 | 6 | 0 | 0
04 Nov 2022 | CH-LAS | TB | 0 | 0 | 0 | 54 | 0 | 0
04 Nov 2022 | CH-FRU | TB | 0 | 0 | 0 | 34 | 0 | 0
02 Nov 2022 | CH-DAS | LK | 0 | 0 | 0 | 66 | 0 | 0
01 Nov 2022 | CH-LAS | LS | 0 | 0 | 0 | 54 | 0 | 0
01 Nov 2022 | CH-OE2 | FT | 0 | 1 | 0 | 11 | 0 | 0 | LGR program not running since 31.10, today Philip restarted it
31 Oct 2022 | CH-CHA | YW | 0 | 0 | 1 | 71 | 0 | 0 | irgashow: H2O shows negative value
31 Oct 2022 | CH-FRU | LH | 0 | 0 | 0 | 34 | 0 | 0
31 Oct 2022 | CH-DAV | IF | 0 | 0 | 0 | 43 | 0 | 1 | One smaller file on 28.10. 01:00
28 Oct 2022 | CH-LAE | AS | 0 | 1 | 1 | 6 | 0 | 0 | There was some interruption in data acquisition on 26th Oct and 27th Oct. The data looks good however.
27 Oct 2022 | CH-AWS | PM | 0 | 0 | 0 | 55 | 0 | 0
27 Oct 2022 | CH-CHA | PM | 0 | 0 | 0 | 75 | 0 | 0
27 Oct 2022 | CH-DAV | PM | 0 | 0 | 0 | 43 | 0 | 0
27 Oct 2022 | CH-DAS | PM | 0 | 0 | 0 | 66 | 0 | 0
27 Oct 2022 | CH-AWS | GA | 0 | 0 | 0 | 55 | 0 | 0
27 Oct 2022 | CH-OE2 | TB | 0 | 0 | 0 | 11 | 0 | 0
26 Oct 2022 | CH-LAE | TB | 0 | 0 | 1 | 6 | 0 | 0 | The system was down during afternoon. There where mains power problem in the region. It was running ok in the evening
26 Oct 2022 | CH-LAS | TB | 0 | 0 | 0 | 54 | 0 | 0 | IRGA Gain factor at 62.5% which is rather high. Maybe only some water
25 Oct 2022 | CH-OE2 | FT | 0 | 0 | 0 | 11 | 0 | 0
24 Oct 2022 | CH-FRU | LH | 0 | 0 | 0 | 34 | 0 | 0
24 Oct 2022 | CH-CHA | YW | 0 | 0 | 0 | 71 | 0 | 0 | Disk use high. Need to delete some files.
24 Oct 2022 | CH-LAS | LS | 0 | 0 | 0 | 53 | 0 | 0
24 Oct 2022 | CH-DAV | IF | 0 | 0 | 0 | 42 | 0 | 0
22 Oct 2022 | CH-AWS | GA | 0 | 0 | 0 | 54 | 0 | 0
19 Oct 2022 | CH-DAS | LK | 0 | 0 | 0 | 65 | 0 | 0
19 Oct 2022 | CH-LAS | LS | 0 | 0 | 0 | 52 | 0 | 0
19 Oct 2022 | CH-OE2 | FT | 0 | 0 | 0 | 11 | 0 | 0 | lgr not working between 16.10 and 17.10
19 Oct 2022 | CH-FRU | LH | 0 | 0 | 1 | 34 | 0 | 0 | Window 93.8% RAIN or NEEDS CLEANING
17 Oct 2022 | CH-CHA | YW | 0 | 0 | 0 | 70 | 0 | 0
17 Oct 2022 | CH-DAV | IF | 0 | 0 | 0 | 41 | 0 | 0
14 Oct 2022 | CH-AWS | GA | 0 | 0 | 0 | 53 | 0 | 0
13 Oct 2022 | CH-LAS | TB | 0 | 0 | 0 | 51 | 0 | 0
13 Oct 2022 | CH-LAE | TB | 0 | 0 | 1 | 6 | 0 | 0
13 Oct 2022 | CH-OE2 | TB | 0 | 0 | 0 | 11 | 0 | 0 | lgr is running and lgrshow works
11 Oct 2022 | CH-DAS | LK | 0 | 0 | 2 | 63 | 0 | 0 | IRGA shows unrealistic values & constantly shows Window 93.8% RAIN or NEEDS CLEANING
11 Oct 2022 | CH-AWS | GA | 0 | 0 | 0 | 52 | 0 | 0
11 Oct 2022 | CH-OE2 | FT | 0 | 0 | 0 | 11 | 0 | 0
10 Oct 2022 | CH-CHA | YW | 0 | 0 | 0 | 69 | 0 | 0
10 Oct 2022 | CH-LAS | LS | 0 | 0 | 0 | 51 | 0 | 0
10 Oct 2022 | CH-FRU | LH | 0 | 0 | 0 | 33 | 0 | 0
10 Oct 2022 | CH-DAV | IF | 0 | 0 | 0 | 41 | 0 | 0
05 Oct 2022 | CH-LAE | AS | 0 | 0 | 1 | 6 | 0 | 0 | Old Irga problem with flow module. See previous checks.
04 Oct 2022 | CH-OE2 | FT | 0 | 0 | 0 | 11 | 0 | 0 | the LGR was shut down between 29.09 to 03.10
04 Oct 2022 | CH-FRU | LH | 0 | 0 | 0 | 33 | 0 | 0
03 Oct 2022 | CH-AWS | PM | 0 | 0 | 0 | 51 | 0 | 0
03 Oct 2022 | CH-DAV | PM | 0 | 0 | 0 | 40 | 0 | 0
03 Oct 2022 | CH-DAS | PM | 0 | 0 | 1 | 62 | 0 | 0 | IRGA: Window 93.8% RAIN or NEEDS CLEANING
03 Oct 2022 | CH-CHA | YW | 0 | 0 | 0 | 68 | 0 | 0
03 Oct 2022 | CH-DAV | IF | 0 | 0 | 0 | 40 | 0 | 0
03 Oct 2022 | CH-LAS | LS | 0 | 0 | 0 | 50 | 0 | 0
30 Sep 2022 | CH-DAS | LK | 0 | 0 | 0 | 62 | 0 | 0
29 Sep 2022 | CH-AWS | GA | 0 | 0 | 0 | 50 | 0 | 0
27 Sep 2022 | CH-OE2 | FT | 0 | 0 | 0 | 10 | 0 | 0 | IRGA signal 93.4%
26 Sep 2022 | CH-DAV | IF | 0 | 0 | 0 | 38 | 0 | 0
26 Sep 2022 | CH-CHA | YW | 0 | 0 | 0 | 67 | 0 | 0
26 Sep 2022 | CH-LAS | LS | 0 | 0 | 0 | 49 | 0 | 0 | No data between 06.09.2022 and 16.09.2022
26 Sep 2022 | CH-FRU | LH | 0 | 0 | 0 | 33 | 0 | 0
24 Sep 2022 | CH-AWS | GA | 0 | 0 | 0 | 48 | 0 | 0
23 Sep 2022 | CH-DAV | PM | 0 | 0 | 0 | 38 | 0 | 0
23 Sep 2022 | CH-DAS | PM | 0 | 0 | 0 | 61 | 0 | 0
23 Sep 2022 | CH-AWS | PM | 0 | 0 | 0 | 49 | 0 | 0
22 Sep 2022 | CH-DAS | LK | 0 | 0 | 0 | 61 | 0 | 0
22 Sep 2022 | CH-LAE | TB | 0 | 0 | 1 | 5 | 0 | 0
22 Sep 2022 | CH-OE2 | TB | 0 | 0 | 0 | 10 | 0 | 0
21 Sep 2022 | CH-DAV | IF | 0 | 0 | 0 | 37 | 0 | 0
20 Sep 2022 | CH-LAS | LS | 0 | 0 | 0 | 48 | 0 | 0
20 Sep 2022 | CH-OE2 | FT | 0 | 0 | 1 | 10 | 0 | 0 | IRGA signal 93.4no data from LGR from 11.09 to 16.09
19 Sep 2022 | CH-FRU | LH | 0 | 0 | 0 | 33 | 0 | 0
19 Sep 2022 | CH-CHA | YW | 0 | 0 | 0 | 66 | 0 | 0
16 Sep 2022 | CH-AWS | GA | 0 | 0 | 0 | 48 | 0 | 0
13 Sep 2022 | CH-OE2 | FT | 0 | 0 | 1 | 10 | 0 | 1 | no LGR data since September 11 afternoon
13 Sep 2022 | CH-LAS | LS | 2 | [5 empty] | No access to ssh "ssh: connect to host subcanopy port 22: No route to host"
13 Sep 2022 | CH-DAS | LK | 0 | 0 | 0 | 60 | 0 | 0
12 Sep 2022 | CH-DAV | IF | 0 | 0 | 0 | 37 | 0 | 0 | Irga flowrate at 13.7
12 Sep 2022 | CH-FRU | LH | 0 | 0 | 0 | 33 | 0 | 0
12 Sep 2022 | CH-CHA | YW | 0 | 0 | 0 | 65 | 0 | 0
08 Sep 2022 | CH-DAS | LK | 0 | 0 | 0 | 59 | 0 | 0
06 Sep 2022 | CH-OE2 | FT | 0 | 0 | 0 | 10 | 0 | 0 | lgr is now running in CH-OE2
05 Sep 2022 | CH-DAS | LK | 0 | 0 | 0 | 59 | 0 | 0
05 Sep 2022 | CH-FRU | LH | 0 | 0 | 0 | 33 | 0 | 0
05 Sep 2022 | CH-CHA | YW | 0 | 0 | 0 | 64 | 0 | 0
05 Sep 2022 | CH-LAS | LS | 0 | 0 | 0 | 47 | 0 | 0
05 Sep 2022 | CH-DAV | IF | 0 | 0 | 0 | 36 | 0 | 0
02 Sep 2022 | CH-AWS | GA | 0 | 0 | 1 | 45 | 0 | 0 | AGC of irga is on 93%
01 Sep 2022 | CH-LAE | TB | 0 | 0 | 1 | 5 | 0 | 0
01 Sep 2022 | CH-OE2 | TB | 0 | 0 | 0 | 10 | 0 | 0
30 Aug 2022 | CH-DAV | IF | 0 | 0 | 0 | 36 | 0 | 1 | Two smaller-than-usual files on 25.8.
30 Aug 2022 | CH-DAS | LK | 0 | 0 | 0 | 58 | 0 | 0
30 Aug 2022 | CH-OE2 | FT | 0 | 0 | 0 | 10 | 0 | 0 | webcam needs cleaning
29 Aug 2022 | CH-DAV | PM | 0 | 0 | 0 | 35 | 0 | 0 | Planned power outage on 22nd of August for about 15 min.
29 Aug 2022 | CH-DAS | PM | 0 | 0 | 0 | 58 | 0 | 0
29 Aug 2022 | CH-AWS | PM | 0 | 0 | 0 | 45 | 0 | 0
29 Aug 2022 | CH-CHA | YW | 0 | 0 | 1 | 63 | 0 | 0 | Irga window needs cleaning.
29 Aug 2022 | CH-FRU | LH | 0 | 0 | 0 | 32 | 0 | 0
26 Aug 2022 | CH-AWS | GA | 0 | 0 | 0 | 44 | 0 | 0
25 Aug 2022 | CH-DAV | IF | 0 | 0 | 0 | 35 | 0 | 2 | file 2022082207.X00 is smaller than usual (29 MB instead of 31 MB). No data in gl-archive since 22.8.2022
25 Aug 2022 | CH-OE2 | TB | 0 | 0 | 0 | 9 | 0 | 0
25 Aug 2022 | CH-OE2 | FT | 0 | 0 | 0 | 9 | 0 | 0
25 Aug 2022 | CH-LAE | TB | 0 | 0 | 1 | 5 | 0 | 0 | no flow module data
25 Aug 2022 | CH-DAS | LK | 0 | 0 | 0 | 57 | 0 | 0
22 Aug 2022 | CH-FRU | LH | 0 | 0 | 0 | 32 | 0 | 0
22 Aug 2022 | CH-DAS | LK | 0 | 1 | 0 | 57 | 0 | 0
22 Aug 2022 | CH-CHA | YW | 0 | 0 | 0 | 62 | 0 | 0
22 Aug 2022 | CH-LAE | AS | 0 | 0 | 1 | 5 | 0 | 0 | IRGA AGC is now 86.7%. Will be cleaned on next tower climb (later this week or next week).The same IRGA issue with flow data is ongoing (see previous entries).
22 Aug 2022 | CH-LAS | LS | 0 | 0 | 0 | 45 | 0 | 0
19 Aug 2022 | CH-LAE | TB | 0 | 0 | 1 | 5 | 0 | 0 | IRGA Flow is at 86.7. We need to observe. Most likely due to some rain
18 Aug 2022 | CH-AWS | GA | 0 | 0 | 1 | 43 | 0 | 0 | irga CO2 values are high
18 Aug 2022 | CH-OE2 | TB | 0 | 0 | 0 | 9 | 0 | 0
17 Aug 2022 | CH-OE2 | FT | 0 | 0 | 0 | 9 | 0 | 0
16 Aug 2022 | CH-CHA | YW | 0 | 0 | 0 | 62 | 0 | 0
15 Aug 2022 | CH-DAV | LH | 0 | 0 | 0 | 34 | 0 | 0
15 Aug 2022 | CH-FRU | LH | 0 | 0 | 0 | 32 | 0 | 0
12 Aug 2022 | CH-LAS | LS | 0 | 0 | 0 | 44 | 0 | 0
12 Aug 2022 | CH-LAE | AS | 0 | 0 | 1 | 5 | 0 | 0 | Issues with irga and flow data. See previous entries.
11 Aug 2022 | CH-DAS | LK | 0 | 0 | 0 | 55 | 0 | 0
11 Aug 2022 | CH-AWS | GA | 0 | 0 | 0 | 42 | 0 | 0
09 Aug 2022 | CH-AWS | PM | 0 | 0 | 0 | 41 | 0 | 0
09 Aug 2022 | CH-DAV | PM | 0 | 0 | 0 | 33 | 0 | 0
09 Aug 2022 | CH-OE2 | FT | 0 | 0 | 0 | 9 | 0 | 0
08 Aug 2022 | CH-CHA | YW | 0 | 0 | 0 | 60 | 0 | 0
08 Aug 2022 | CH-FRU | LH | 0 | 0 | 0 | 32 | 0 | 0
08 Aug 2022 | CH-DAV | LH | 0 | 0 | 0 | 33 | 0 | 0
05 Aug 2022 | CH-DAV | LH | 0 | 0 | 0 | 32 | 0 | 0
05 Aug 2022 | CH-FRU | LH | 0 | 0 | 0 | 32 | 0 | 0
05 Aug 2022 | CH-AWS | GA | 0 | 0 | 0 | 40 | 0 | 0
04 Aug 2022 | CH-OE2 | TB | 0 | 0 | 0 | 9 | 0 | 0
04 Aug 2022 | CH-LAE | TB | 0 | 0 | 0 | 5 | 0 | 0
04 Aug 2022 | CH-DAS | LK | 0 | 0 | 0 | 54 | 0 | 0
02 Aug 2022 | CH-CHA | YW | 0 | 0 | 0 | 60 | 0 | 0
02 Aug 2022 | CH-LAE | AS | 0 | 0 | 1 | 4 | 0 | 0 | Old Irga problem with flow module. Please see previous entries
02 Aug 2022 | CH-OE2 | FT | 0 | 0 | 0 | 9 | 0 | 0
02 Aug 2022 | CH-DAS | LK | 0 | 2 | 0 | 54 | 2 | 0 | sonic again stopped working on 2.8.
02 Aug 2022 | CH-LAS | AS | 0 | 0 | 0 | 40 | 0 | 0
02 Aug 2022 | CH-DAV | LH | 0 | 0 | 0 | 32 | 0 | 0
02 Aug 2022 | CH-FRU | LH | 0 | 0 | 0 | 32 | 0 | 0
28 Jul 2022 | CH-DAV | LH | 0 | 0 | 0 | 31 | 0 | 0
28 Jul 2022 | CH-FRU | LH | 0 | 0 | 0 | 32 | 0 | 0
28 Jul 2022 | CH-AWS | GA | 0 | 0 | 0 | 39 | 0 | 0
25 Jul 2022 | CH-DAS | LK | 0 | 2 | 0 | 53 | 2 | 2 | no files since the 23.07. sonic seems to be not working
25 Jul 2022 | CH-CHA | YW | 0 | 0 | 0 | 59 | 0 | 0
25 Jul 2022 | CH-DAV | LH | 0 | 0 | 0 | 31 | 0 | 0
25 Jul 2022 | CH-FRU | LH | 0 | 0 | 0 | 32 | 0 | 0
25 Jul 2022 | CH-LAE | AS | 0 | 0 | 1 | 4 | 0 | 0 | Old irga issue with flowrate. The AGC is 93.4% (since last week).
25 Jul 2022 | CH-LAS | LS | 0 | 0 | 0 | 41 | 0 | 0
25 Jul 2022 | CH-OE2 | RM | 0 | 0 | 0 | 9 | 0 | 0 | AGC 93.4%
21 Jul 2022 | CH-AWS | GA | 0 | 0 | 0 | 38 | 0 | 0
18 Jul 2022 | CH-CHA | YW | 0 | 0 | 0 | 58 | 0 | 0
18 Jul 2022 | CH-FRU | LH | 0 | 0 | 0 | 31 | 0 | 0
18 Jul 2022 | CH-DAV | LH | 0 | 0 | 0 | 30 | 0 | 0
18 Jul 2022 | CH-OE2 | FT | 0 | 0 | 0 | 9 | 0 | 0
17 Jul 2022 | CH-DAV | LH | 0 | 0 | 0 | 30 | 0 | 0
17 Jul 2022 | CH-FRU | LH | 0 | 0 | 0 | 31 | 0 | 0
16 Jul 2022 | CH-AWS | GA | 0 | 0 | 0 | 36 | 0 | 0
14 Jul 2022 | CH-DAS | LK | 0 | 0 | 0 | 52 | 0 | 0
12 Jul 2022 | CH-OE2 | FT | 0 | 0 | 0 | 9 | 0 | 0 | AGC 93.4%
11 Jul 2022 | CH-FRU | LH | 0 | 0 | 0 | 31 | 0 | 0 | Level-0 fluxes look OK
11 Jul 2022 | CH-DAV | LH | 0 | 0 | 0 | 29 | 0 | 0 | Level-0 fluxes look OK
11 Jul 2022 | CH-CHA | YW | 0 | 0 | 0 | 57 | 0 | 0 | No LGR data since July 6 due to instrument removal from the site. File size is therefore reduced by half.
10 Jul 2022 | CH-AWS | GA | 0 | 0 | 0 | 36 | 0 | 0
07 Jul 2022 | CH-DAS | LK | 0 | 0 | 0 | 51 | 0 | 0
07 Jul 2022 | CH-LAE | AS | 0 | 0 | 1 | 4 | 0 | 0 | Same flowrate issue with irgashow (please see previous entries)
05 Jul 2022 | CH-OE2 | FT | 0 | 0 | 0 | 8 | 0 | 0 | AGC 93.4%
05 Jul 2022 | CH-FRU | LH | 0 | 0 | 0 | 31 | 0 | 0 | Level-0 fluxes also look good
05 Jul 2022 | CH-DAV | LH | 0 | 0 | 0 | 29 | 0 | 0 | Level-0 fluxes also look good
04 Jul 2022 | CH-CHA | YW | 0 | 0 | 0 | 55 | 1 | 0 | There are 3 files smaller than 26 MB for Jun 30, and one for July 2. Will check what is missing.
02 Jul 2022 | CH-AWS | GA | 0 | 0 | 0 | 94 | 0 | 0 | irga AGC again on 100%
29 Jun 2022 | CH-DAS | LK | 0 | 0 | 0 | 50 | 0 | 0
29 Jun 2022 | CH-OE2 | FT - RM | 0 | 0 | 0 | 8 | 0 | 0 | AGC 93.4%
28 Jun 2022 | CH-AWS | PM | 0 | 0 | 1 | 93 | 0 | 0 | IRGA AGC at 87%, needs cleaning
28 Jun 2022 | CH-DAV | PM | 0 | 0 | 1 | 28 | 0 | 0 | IRGA AGC at 87%, needs cleaning
27 Jun 2022 | CH-CHA | YW | 0 | 0 | 0 | 54 | 0 | 0
27 Jun 2022 | CH-FRU | LH | 0 | 0 | 0 | 31 | 0 | 0
27 Jun 2022 | CH-DAV | LH | 0 | 0 | 0 | 28 | 0 | 0
24 Jun 2022 | CH-LAE | AS | 0 | 1 | 1 | 4 | 0 | 0 | Old Flow module problem.
24 Jun 2022 | CH-LAS | LS | 0 | 0 | 0 | 37 | 0 | 0
23 Jun 2022 | CH-DAS | LK | 0 | 0 | 1 | 49 | 0 | 0 | IRGA window dirtiness at 56.2%
22 Jun 2022 | CH-OE2 | FT | 0 | 0 | 0 | 8 | 0 | 0
22 Jun 2022 | CH-AWS | GA | 0 | 0 | 1 | 92 | 0 | 0 | on Irgashow CO2 around 450
21 Jun 2022 | CH-FRU | LH | 0 | 0 | 0 | 31 | 0 | 0
21 Jun 2022 | CH-DAV | LH | 0 | 0 | 0 | 27 | 0 | 0
20 Jun 2022 | CH-DAS | LK | 0 | 0 | 0 | 49 | 0 | 0
20 Jun 2022 | CH-CHA | YW | 0 | 0 | 0 | 52 | 0 | 0 | irga window slightly dirty (75%). Cleaning needed.
15 Jun 2022 | CH-OE2 | RM/FT | 0 | 0 | 0 | 8 | 0 | 0
14 Jun 2022 | CH-LAS | LS | 0 | 0 | 0 | 35 | 0 | 0
14 Jun 2022 | CH-LAE | AS | 0 | 0 | 1 | 4 | 0 | 0 | Old Irga problem with flow module
14 Jun 2022 | CH-DAS | LK | 0 | 0 | 0 | 48 | 0 | 0
13 Jun 2022 | CH-CHA | YW | 0 | 0 | 0 | 50 | 0 | 0
13 Jun 2022 | CH-FRU | LH | 0 | 0 | 1 | 31 | 0 | 0 | Window 93.8% RAIN or NEEDS CLEANING
13 Jun 2022 | CH-DAV | LH | 0 | 0 | 0 | 26 | 0 | 0
13 Jun 2022 | CH-AWS | GA | 0 | 0 | 0 | 91 | 0 | 1 | Snowhight jumping between 10^4 Order of magnitudes
10 Jun 2022 | CH-OE2 | TB | 0 | 0 | 0 | 8 | 0 | 0
10 Jun 2022 | CH-LAE | TB | 0 | 0 | 1 | 4 | 0 | 0 | still no flow module data
10 Jun 2022 | CH-LAS | TB | 0 | 0 | 0 | 35 | 0 | 0
10 Jun 2022 | CH-OE2 | RM | 0 | 0 | 0 | 8 | 0 | 0
09 Jun 2022 | CH-LAS | LS | 0 | 0 | 1 | 35 | 0 | 0 | Rain or needs to be clean
09 Jun 2022 | CH-DAS | LK | 0 | 0 | 0 | 47 | 0 | 0
08 Jun 2022 | CH-CHA | YW | 0 | 0 | 0 | 49 | 0 | 0 | Everything back to normal. .C00 file size was smaller (without LGR data) from June 4 to June 7.
07 Jun 2022 | CH-CHA | YW | 0 | 1 | 0 | 48 | 1 | 1 | No LGR Data and .C00 file size reduced by half since June 5. Level-0 fluxes look OK for data before June 3.
07 Jun 2022 | CH-AWS | GA | 0 | 0 | 1 | 90 | 0 | 0 | on irgashow CO2 around 450
06 Jun 2022 | CH-FRU | LH | 0 | 0 | 0 | 30 | 0 | 0
06 Jun 2022 | CH-DAV | LH | 0 | 0 | 0 | 25 | 0 | 0
02 Jun 2022 | CH-DAS | LK | 0 | 0 | 0 | 46 | 0 | 0
01 Jun 2022 | CH-CHA | MS | 0 | 0 | 0 | 50 | 0 | 0
31 May 2022 | CH-FRU | LH | 0 | 0 | 0 | 30 | 0 | 0 | Level-0 fluxes also look OK
31 May 2022 | CH-DAV | LH | 0 | 0 | 0 | 24 | 0 | 0 | Level-0 fluxes also look OK
31 May 2022 | CH-LAS | LS | 0 | 0 | 0 | 33 | 0 | 0
31 May 2022 | CH-OE2 | RM | 0 | 0 | 0 | 8 | 0 | 0
30 May 2022 | CH-AWS | GA | 0 | 0 | 0 | 88 | 0 | 0
30 May 2022 | CH-CHA | YW | 0 | 0 | 0 | 47 | 0 | 0
26 May 2022 | CH-FRU | LH | 0 | 0 | 0 | 30 | 0 | 0
26 May 2022 | CH-DAV | LH | 0 | 0 | 0 | 24 | 0 | 0
24 May 2022 | CH-CHA | YW | 0 | 0 | 0 | 45 | 0 | 0
24 May 2022 | CH-OE2 | RM | 0 | 0 | 0 | 8 | 0 | 0
23 May 2022 | CH-DAS | LK | 0 | 0 | 0 | 45 | 0 | 0
20 May 2022 | CH-AWS | GA | 0 | 0 | 0 | 87 | 0 | 0
19 May 2022 | CH-CHA | MS | 0 | 0 | 0 | 90 | 1 | 0 | On may 12 and 13 single files are smaller, no idea why.
19 May 2022 | CH-FRU | MS | 0 | 0 | 0 | 30 | 0 | 0
19 May 2022 | CH-LAE | TB | 0 | 0 | 0 | 3 | 0 | 0
19 May 2022 | CH-OE2 | TB | 0 | 0 | 0 | 8 | 0 | 0
18 May 2022 | CH-CHA | YW | 0 | 0 | 0 | 85 | 0 | 0
17 May 2022 | CH-LAS | LS | 0 | 0 | 0 | 31 | 0 | 0
17 May 2022 | CH-DAV | PM | 0 | 0 | 1 | 23 | 0 | 0 | AGC at 86.7 %. Needs cleaning at the next site visit
17 May 2022 | CH-DAS | PM | 0 | 0 | 0 | 44 | 0 | 0
17 May 2022 | CH-AWS | PM | 0 | 0 | 0 | 86 | 0 | 0 | switched off mast height control, winter is over
16 May 2022 | CH-FRU | LH | 0 | 0 | 0 | 30 | 0 | 0
16 May 2022 | CH-DAV | LH | 0 | 0 | 0 | 22 | 0 | 0 | signal strength slightly reduced at 86.7%
16 May 2022 | CH-LAE | AS | 0 | 1 | 1 | 3 | 0 | 0 | Old flow module problem
13 May 2022 | CH-LAE | TB | 0 | 1 | 1 | 3 | 0 | 0 | AGC is 100% again. Flow module --> not data transfer to analyzer
13 May 2022 | CH-OE2 | TB | 0 | 0 | 0 | 7 | 0 | 0 | AGC is 100% again
13 May 2022 | CH-LAS | TB | 0 | 0 | 0 | 31 | 0 | 0
12 May 2022 | CH-LAS | LS | 0 | 0 | 0 | 31 | 0 | 0
12 May 2022 | CH-DAS | LK | 0 | 0 | 0 | 43 | 0 | 0
12 May 2022 | CH-AWS | GA | 0 | 0 | 0 | 85 | 0 | 0
10 May 2022 | CH-CHA | YW | 0 | 0 | 0 | 83 | 0 | 0
09 May 2022 | CH-FRU | LH | 0 | 0 | 0 | 30 | 0 | 0
09 May 2022 | CH-DAV | LH | 0 | 0 | 0 | 22 | 0 | 0
06 May 2022 | CH-DAS | LK | 0 | 0 | 0 | 43 | 0 | 0
06 May 2022 | CH-AWS | GA | 0 | 0 | 0 | 84 | 0 | 0
04 May 2022 | CH-CHA | YW | 0 | 0 | 0 | 81 | 0 | 0 | LGR pressure back to 45 Torr (filter changed 03 May 2022)
03 May 2022 | CH-DAS | LK | 0 | 0 | 0 | 42 | 0 | 0
03 May 2022 | CH-LAS | LS | 0 | 0 | 0 | 29 | 0 | 0
03 May 2022 | CH-DAV | PM | 0 | 0 | 1 | 21 | 0 | 0 | IRGA AGC at 86.7%, needs cleaning
03 May 2022 | CH-AWS | PM | 0 | 0 | 0 | 84 | 0 | 0
30 Apr 2022 | CH-FRU | LH | 0 | 0 | 0 | 29 | 0 | 0
30 Apr 2022 | CH-DAV | LH | 0 | 0 | 0 | 21 | 0 | 0
29 Apr 2022 | CH-AWS | GA | 0 | 0 | 0 | 83 | 0 | 0
28 Apr 2022 | CH-LAE | TB | 0 | 1 | 1 | 3 | 0 | 0
28 Apr 2022 | CH-OE2 | TB | 0 | 0 | 0 | 7 | 0 | 0
28 Apr 2022 | CH-LAS | TB | 0 | 0 | 0 | 28 | 0 | 0
27 Apr 2022 | CH-OE2 | RM | 0 | 0 | 0 | 7 | 0 | 0 | AGC 86.7%, IRGA needs cleaning soon
27 Apr 2022 | CH-CHA | YW | 0 | 0 | 0 | 79 | 0 | 0 | LGR pressure at 27.9 Torr
25 Apr 2022 | CH-FRU | MS | 0 | 0 | 0 | 29 | 0 | 0
25 Apr 2022 | CH-DAV | PM | 0 | 0 | 0 | 20 | 0 | 0
25 Apr 2022 | CH-DAS | PM | 0 | 0 | 0 | 41 | 0 | 0
25 Apr 2022 | CH-AWS | PM | 0 | 2 | 0 | 82 | 1 | 1 | sonic not able to measure wind speeds due to snowfall
22 Apr 2022 | CH-AWS | GA | 0 | 0 | 0 | 82 | 0 | 0
21 Apr 2022 | CH-DAS | LK | 0 | 0 | 0 | 41 | 0 | 0
21 Apr 2022 | CH-LAE | TB | 0 | 1 | 1 | 3 | 0 | 0 | AGC again at 86.7%. We should clean soon
20 Apr 2022 | CH-LAE | AS | 0 | 1 | 1 | 3 | 0 | 0 | The existing problem with zero flow rate message
20 Apr 2022 | CH-FRU | AS | 0 | 0 | 0 | 29 | 0 | 0 | All good so far.
20 Apr 2022 | CH-CHA | YW | 0 | 0 | 0 | 77 | 0 | 0 | 32.7 Torr LGR pressure
19 Apr 2022 | CH-LAS | LS | 0 | 0 | 0 | 27 | 0 | 0
15 Apr 2022 | CH-LAE | TB | 0 | 1 | 1 | 3 | 0 | 0
14 Apr 2022 | CH-DAS | LK | 0 | 0 | 0 | 40 | 0 | 0
14 Apr 2022 | CH-AWS | GA | 0 | 0 | 0 | 81 | 0 | 0
12 Apr 2022 | CH-DAV | PM | 0 | 0 | 0 | 18 | 0 | 0
12 Apr 2022 | CH-AWS | PM | 0 | 0 | 0 | 80 | 0 | 0
12 Apr 2022 | CH-CHA | YW | 0 | 0 | 0 | 75 | 0 | 0 | LGR pressure at 31 Torr
11 Apr 2022 | CH-FRU | LH | 0 | 0 | 0 | 29 | 0 | 0 | Level-0 fluxes also look good
08 Apr 2022 | CH-LAE | TB | 0 | 1 | 0 | 2 | 0 | 0
07 Apr 2022 | CH-AWS | GA | 0 | 0 | 0 | 79 | 0 | 0
05 Apr 2022 | CH-DAV | PM | 0 | 0 | 0 | 17 | 0 | 0
05 Apr 2022 | CH-DAS | PM | 0 | 0 | 0 | 38 | 0 | 0
04 Apr 2022 | CH-FRU | LH | 0 | 0 | 0 | 29 | 0 | 1 | only sporadic files between 31 Mar and 1 Apr, but system running file since 2 Apr morning
04 Apr 2022 | CH-DAV | LH | 0 | 0 | 0 | 17 | 0 | 0
04 Apr 2022 | CH-LAS | LS | 0 | 0 | 0 | 25 | 0 | 0 | Window dirtiness more than 60%
04 Apr 2022 | CH-AWS | PM | 0 | 0 | 0 | 79 | 0 | 0 | Mast went up to 3.8 m (!) on April 3rd and to 2.5 m (!) today around noon. Check GIN for details
01 Apr 2022 | CH-AWS | GA | 0 | 0 | 0 | 78 | 0 | 0
30 Mar 2022 | CH-CHA | IF | 0 | 0 | 0 | 71 | 0 | 0 | LGR pressure at 35 Torr, one week after filter change
28 Mar 2022 | CH-FRU | LH | 0 | 0 | 0 | 29 | 0 | 0
28 Mar 2022 | CH-DAV | LH | 0 | 0 | 0 | 16 | 0 | 0
28 Mar 2022 | CH-DAV | PM | 0 | 0 | 0 | 16 | 0 | 0
28 Mar 2022 | CH-DAS | PM | 0 | 0 | 0 | 37 | 0 | 0
28 Mar 2022 | CH-AWS | PM | 0 | 1 | 2 | 78 | 1 | 0 | No IRGA data since March 23. Thus, smaller file sizes.
25 Mar 2022 | CH-OE2 | RM | 0 | 0 | 0 | 6 | 0 | 0
25 Mar 2022 | CH-LAE | TB | 0 | 1 | 1 | 2 | 0 | 0
25 Mar 2022 | CH-OE2 | TB | 0 | 0 | 0 | 6 | 0 | 0
24 Mar 2022 | CH-LAE | AS | 0 | 1 | 1 | 2 | 0 | 0 | Previous problem with IRGA data still existing. But the calculated fluxes look okay
23 Mar 2022 | CH-FRU | LH | 0 | 0 | 0 | 29 | 0 | 0 | recent Level-0 fluxes also look good
23 Mar 2022 | CH-DAV | LH | 0 | 0 | 0 | 16 | 0 | 0 | recent Level-0 fluxes also look good
23 Mar 2022 | CH-CHA | IF | 0 | 0 | 0 | 70 | 0 | 0 | After filter change LGR pressure is at 45 Torr again
21 Mar 2022 | CH-AWS | PM | 0 | 0 | 0 | 77 | 0 | 0
21 Mar 2022 | CH-DAV | PM | 0 | 0 | 0 | 15 | 0 | 0
21 Mar 2022 | CH-DAS | PM | 0 | 0 | 0 | 36 | 0 | 0
21 Mar 2022 | CH-LAS | LS | 0 | 0 | 0 | 23 | 0 | 0
18 Mar 2022 | CH-LAE | TB | 0 | 1 | 1 | 2 | 0 | 0
18 Mar 2022 | CH-OE2 | TB | 0 | 0 | 0 | 6 | 0 | 0
16 Mar 2022 | CH-OE2 | RM | 0 | 0 | 0 | 6 | 0 |  | Cannot check if files are here but everything looks good though. AGC at 93.4%.
15 Mar 2022 | CH-CHa | MS | 0 | 0 | 0 | 72 | 0 | 0
15 Mar 2022 | CH-FRU | MS | 0 | 0 | 0 | 28 | 0 | 0
14 Mar 2022 | CH-DAV | LH | 0 | 0 | 0 | 15 | 0 | 0
14 Mar 2022 | CH-FRU | LH | 0 | 0 | 0 | 28 | 0 | 0
14 Mar 2022 | CH-LAE | AS | 0 | 1 | 1 | 2 | 1 | 0 | The flow rate in IRGA data file is zero but flow module on the tower is showing the desired flowrate of 12. So maybe issue with the cable. Thomas is checking. Also, the sonicshow is showing NO IRGA DATA every 10 seconds (Thomas changed the buffer to 10 seconds in the sonic read script on Friday 11th March 2022). All the other IRGA data are fine.
14 Mar 2022 | CH-AWS | PM | 0 | 0 | 0 | 76 | 0 | 0 | QA/QC from the other side of the world works seemlessly :-)
14 Mar 2022 | CH-DAV | PM | 0 | 0 | 0 | 15 | 0 | 0
14 Mar 2022 | CH-DAS | PM | 0 | 0 | 0 | 35 | 0 | 0
11 Mar 2022 | CH-LAS | LS | 0 |  | 0 | 22 | 0 | 0
11 Mar 2022 | CH-AWS | GA | 0 | 0 | 0 | 75 | 0 | 0
10 Mar 2022 | CH-DAS | LK | 0 | 0 | 0 | 35 | 0 | 0
10 Mar 2022 | CH-OE2 | RM | 0 | 0 | 0 | 6 | 0 | 0
08 Mar 2022 | CH-LAE | TB | 0 | 0 | 2 | 2 | 0 | 0 | LI7200 flow is zero
08 Mar 2022 | CH-OE2 | TB | 0 | 0 | 0 | 6 | 0 | 0
07 Mar 2022 | CH-DAV | PM | 0 | 0 | 0 | 14 | 0 | 0
07 Mar 2022 | CH-AWS | PM | 0 | 0 | 0 | 75 | 0 | 0
07 Mar 2022 | CH-AWS | PM | 0 | 0 | 0 | 75 | 0 | 0 | Mast occasionally frozen, does not always adapt to snowheight. Natural or induced heating solves the problem.
07 Mar 2022 | CH-FRU | LH | 0 | 0 | 0 | 28 | 0 | 0 | Level-0 fluxes also look good
04 Mar 2022 | CH-LAE | AS | 0 | 0 | 0 | 2 | 0 | 0 | All Good
04 Mar 2022 | CH-DAS | LK | 0 | 0 | 0 | 34 | 0 | 0
03 Mar 2022 | CH-AWS | GA | 0 | 0 | 0 | 74 | 0 | 0
03 Mar 2022 | CH-CHA | IF | 0 | 0 | 1 | 65 | 0 | 0 | 29 Torr cell pressure, CH4 concentration is quite high (>3 ppm)
02 Mar 2022 | CH-OE2 | RM | 0 | 0 | 0 | 6 | 0 | 0
02 Mar 2022 | CH-LAS | LS | 0 | 0 | 0 | 21 | 0 | 0
01 Mar 2022 | CH-DAV | PM | 0 | 0 | 0 | 13 | 0 | 0
01 Mar 2022 | CH-AWS | PM | 0 | 0 | 0 | 74 | 0 | 0 | Mast was not moving. Probably frozen.
28 Feb 2022 | CH-FRU | LH | 0 | 0 | 0 | 28 | 0 | 0 | Level-0 fluxes also look OK
28 Feb 2022 | CH-DAV | LH | 0 | 0 | 0 | 13 | 0 | 0 | Level-0 fluxes also look OK
28 Feb 2022 | CH-CHA | IF | 0 | 0 | 0 | 64 | 0 | 0 | LGR pressure at 35.2 Torr. Sonic temperature jumps of ~1 °C
25 Feb 2022 | CH-LAS | TB | 0 | 0 | 0 | 37 | 0 | 0
25 Feb 2022 | CH-LAE | TB | 0 | 0 | 0 | 2 | 0 | 0
25 Feb 2022 | CH-AWS | GA | 0 | 0 | 0 | 73 | 0 | 0
24 Feb 2022 | CH-DAS | LK | 0 | 0 | 0 | 33 | 0 | 0
24 Feb 2022 | CH-OE2 | RM | 0 | 0 | 0 | 6 | 0 | 0
24 Feb 2022 | CH-OE2 | TB | 0 | 0 | 0 | 6 | 0 | 0
22 Feb 2022 | CH-LAS | LS | 0 | 0 | 0 | 19 | 0 | 0
22 Feb 2022 | CH-CHA | MS | 0 | 0 | 0 | 67 | 0 | 0 | lgr pressure at around 35 torr
22 Feb 2022 | CH-LAE | AS | 0 | 0 | 0 | 2 | 0 | 0 | Irga and sonic both running fine now. The irga was brought to ETH for recalibration on 15.02.2022 and mounted on 18.02.2022.
21 Feb 2022 | CH-FRU | LH | 0 | 0 | 0 | 28 | 0 | 0 | Level-0 fluxes also look good
21 Feb 2022 | CH-DAV | LH | 0 | 0 | 0 | 12 | 0 | 0 | Level-0 fluxes also look good
17 Feb 2022 | CH-DAV | PM | 0 | 0 | 0 | 11 | 0 | 0
17 Feb 2022 | CH-AWS | PM | 0 | 0 | 0 | 72 | 0 | 0
17 Feb 2022 | CH-OE2 | RM | 0 | 0 | 0 | 6 | 0 | 0
17 Feb 2022 | CH-CHA | IF | 0 | 0 | 0 | 61 | 0 | 0 | LGR pressure at 35.3 Torr
14 Feb 2022 | CH-FRU | LH | 0 | 0 | 0 | 28 | 0 | 0 | Level-0 fluxes also look good
14 Feb 2022 | CH-DAV | LH | 0 | 0 | 0 | 11 | 0 | 0 | filesizes for EC ICOS files and EC binaries look good, Level-0 fluxes look good
14 Feb 2022 | CH-LAS | LS | 0 | 0 | 0 | 18 | 0 | 0
11 Feb 2022 | CH-DAS | LK | 0 | 0 | 0 | 31 | 0 | 0
11 Feb 2022 | CH-OE2 | RM | 0 | 0 | 0 | 6 | 0 | 0 | Some smaller files on 6.2. but since then everything complete
10 Feb 2022 | CH-CHA | IF | 0 | 0 | 0 | 59 | 0 | 1 | Some smaller files from 6.2. to 7.2 (no LGR data?). LGR cell pressure at 41.3 Torr.
10 Feb 2022 | CH-LAE | AS | 0 | 0 | 1 | 2 | 0 | 1 | The filesize are back to the usual 16 MB since 7th Feb and flux calculation also worked well. 4 days of files (2-7th Feb) are missing due to sonic not working. The IRGA AGC increased from 66.7% to 80%, but IRGA needs cleaning. Will be done next week when Thomas is back.
08 Feb 2022 | CH-LAE | MS | 0 | 0 | 1 | 2 | 1 | 2 | AGC 80%No sonic data from 1. - 7. Feb. 2022Prop. corrupt ec data from dec. 14. - feb. 07
08 Feb 2022 | CH-CHA | MS | 0 | 0 | 0 | 59 | 0 | 0
07 Feb 2022 | CH-AWS | PM | 0 | 0 | 0 | 70 | 0 | 0
07 Feb 2022 | CH-DAV | PM | 0 | 0 | 1 | 10 | 0 | 0 | AGC at 86.7%, needs cleaning
07 Feb 2022 | CH-FRU | LH | 0 | 0 | 0 | 28 | 0 | 0
03 Feb 2022 | CH-OE2 | RM | 0 | 1 | 0 | 5 | 0 | 0 | Sonic temperature is jumping again, irga data looks fine though. Let's hope the transducers are fine.
03 Feb 2022 | CH-LAE | AS | 0 | 1 | 1 | 0 | 1 | 1 | Irga AGC still at 66.7% and sonicshow is not showing any data, probably sonic is not working. The filesize is also less. Problem seems to have occured yesterday or today morning.
03 Feb 2022 | CH-AWS | GA | 0 | 0 | 0 | 69 | 0 | 0
01 Feb 2022 | CH-CHA | IF | 0 | 0 | 0 | 57 | 0 | 1 | some smaller files on 28.01. maybe snowing. LGR pressure at 29.5 Torr
01 Feb 2022 | CH-LAE | AS | 0 | 0 | 1 | 0 | 0 | 0 | IRGA AGC is 66.7%. Could be because it is snowing. The flow seems okay.
01 Feb 2022 | CH-LAS | LS | 0 | 0 | 1 | 0 | 0 | 0 | IRGA seems slightly dirty. Might be because of snow. Will clean it next time I go to LAE
01 Feb 2022 | CH-DAS | LK | 0 | 0 | 0 | 30 | 0 | 0
31 Jan 2022 | CH-FRU | LH | 0 | 0 | 0 | 27 | 0 | 0
31 Jan 2022 | CH-DAV | MG | 0 | 0 | 0 | 9 | 0 | 0
28 Jan 2022 | CH-AWS | GA | 0 | 0 | 0 | 68 | 0 | 0
28 Jan 2022 | CH-LAE | TB | 0 | 0 | 0 | 2 | 0 | 0
28 Jan 2022 | CH-OE2 | TB | 0 | 0 | 0 | 5 | 0 | 0
27 Jan 2022 | CH-OE2 | RM | 0 | 0 | 0 | 5 | 0 | 0
27 Jan 2022 | CH-DAS | LK | 0 | 0 | 0 | 29 | 0 | 0
26 Jan 2022 | CH-CHA | IF | 0 | 0 | 0 | 56 | 0 | 0 | LGR pressure at 30 Torr
26 Jan 2022 | CH-DAV | MG | 0 | 0 | 0 | 9 | 0 | 0
24 Jan 2022 | CH-FRU | LH | 0 | 0 | 1 | 27 | 0 | 0 | Window 93.8% RAIN or NEEDS CLEANING
21 Jan 2022 | CH-CHA | IF | 0 | 0 | 0 | 55 | 0 | 0 | LGR pressure at 35 Torr.
21 Jan 2022 | CH-DAS | PM | 0 | 0 | 0 | 28 | 0 | 0
21 Jan 2022 | CH-OE2 | TB | 0 | 0 | 0 | 5 | 0 | 0
21 Jan 2022 | CH-LAE | TB | 0 | 0 | 0 | 1 | 0 | 0
20 Jan 2022 | CH-OE2 | RM | 0 | 0 | 0 | 5 | 0 | 0
20 Jan 2022 | CH-DAS | LK | 0 | 0 | 0 | 28 | 0 | 0
20 Jan 2022 | CH-DAV | PM | 0 | 0 | 0 | 8 | 0 | 0
20 Jan 2022 | CH-AWS | PM | 0 | 0 | 0 | 67 | 0 | 0
20 Jan 2022 | CH-CHA | MS | 0 | 0 | 0 | 54 | 0 | 0
17 Jan 2022 | CH-LAS | AS | 0 | 0 | 0 | 13 | 0 | 0
17 Jan 2022 | CH-LAE | AS | 0 | 0 | 0 | 1 | 0 | 0 | All good
17 Jan 2022 | CH-DAV | MG | 0 | 0 | 0 | 8 | 0 | 0
17 Jan 2022 | CH-FRU | LH | 0 | 0 | 0 | 27 | 0 | 0
14 Jan 2022 | CH-OE2 | TB | 0 | 0 | 0 | 5 | 0 | 0 | online again
14 Jan 2022 | CH-LAE | TB | 0 | 0 | 0 | 1 | 0 | 0
13 Jan 2022 | CH-OE2 | RM | 0 | 0 | 0 | 5 | 0 | 1 | Everything is working fine again. Files were transferred during the night. Only two days with smaller files because connection to IRGA was lost.
13 Jan 2022 | CH-DAS | LK | 0 | 0 | 0 | 27 | 0 | 0
13 Jan 2022 | CH-DAV | PM | 0 | 0 | 0 | 7 | 0 | 0
13 Jan 2022 | CH-DAS | PM | 0 | 0 | 0 | 27 | 0 | 0
13 Jan 2022 | CH-AWS | PM | 0 | 0 | 0 | 66 | 0 | 0
12 Jan 2022 | CH-CHA | IF | 0 | 0 | 0 | 52 | 0 | 1 | Internet was down, now working again. C00 Files still need to be transferred to ETH (will probably happen tonight)
10 Jan 2022 | CH-FRU | LH | 0 | 0 | 0 | 27 | 0 | 1 | data gap on 5 and 6 Jan, running again after power cycle reset
07 Jan 2022 | CH-CHA | IF | 2 | 2 | 2 |  | 2 | 2 | No connection to computer possible. Might be only the internet (modem) (hopefully) or computer is not running.
07 Jan 2022 | CH-OE2 | RM | 2 | 2 | 2 |  | 2 | 2 | Same problem as CHA
06 Jan 2022 | CH-DAS | LK | 0 | 0 | 0 | 26 | 0 | 0
06 Jan 2022 | CH-AWS | GA | 0 | 0 | 0 | 64 | 0 | 0
05 Jan 2022 | CH-FRU | LH | 0 | 2 | 0 | 27 | 2 | 2 | data gap since last check, sonic not running at the moment
03 Jan 2022 | CH-DAV | MG | 0 | 2 | 0 | 6 | 2 | 2 | Data gap since 29.12 (sonic failed and had to be restarted today using the iologik on the tower)
03 Jan 2022 | CH-CHA | IF | 0 | 0 | 0 | 50 | 0 | 0 | Deleted files on Lanner (Webcam and .C00) to free storage. LGR pressure at 45.4 Torr
01 Jan 2022 | CH-FRU | LH | 0 | 0 | 0 | 27 | 0 | 0
30 Dec 2021 | CH-DAS | LK | 0 | 0 | 1 | 25 | 0 | 0 | Irga at 68.8% window dirtiness
29 Dec 2021 | CH-OE2 | RM | 0 | 0 | 0 | 5 | 0 | 0
29 Dec 2021 | CH-DAV | PM | 0 | 0 | 0 | 6 | 0 | 0
29 Dec 2021 | CH-AWS | PM | 0 | 0 | 0 | 63 | 0 | 0
28 Dec 2021 | CH-FRU | MS | 0 | 0 | 0 | 27 | 0 | 0
28 Dec 2021 | CH-CHA | MS | 0 | 0 | 0 | 90 | 0 | 0 | /home is using 90.4%. There will be some old OK.2018 etc. folder deleted in the next couple of days...
27 Dec 2021 | CH-DAV | MG | 0 | 0 | 0 | 6 | 0 | 0
23 Dec 2021 | CH-DAS | LK | 0 | 0 | 0 | 24 | 0 | 0
23 Dec 2021 | CH-LAE | TB | 0 | 0 | 0 | 0 | 0 | 0
23 Dec 2021 | CH-LAS | TB | 1 | 0 | 0 | 0 | 0 | 0 | Direct RSA Key connect not yet established but WE will fix it before christmas
22 Dec 2021 | CH-OE2 | RM | 0 | 0 | 0 | 5 | 0 | 0
21 Dec 2021 | CH-FRU | LH | 0 | 0 | 0 | 26 | 0 | 0
21 Dec 2021 | CH-CHA | IF | 0 | 2 | 0 |  | 1 | 0 | New file created today at 11:00, no GIN entry. Sonic has some problems, high fluctuations in w and sonic temperature since 10th December.
21 Dec 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
20 Dec 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
20 Dec 2021 | CH-FRU | MS | 0 | 0 | 0 |  | 0 | 0
20 Dec 2021 | CH-CHA | MS | 0 | 0 | 0 |  | 0 | 0
18 Dec 2021 | CH-AWS | GA | 0 | 0 | 0 |  | 0 | 0
16 Dec 2021 | CH-DAS | LK | 0 | 0 | 0 |  | 0 | 0
16 Dec 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
16 Dec 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
15 Dec 2021 | CH-OE2 | RM | 0 | 0 | 2 |  | 0 | 0 | IRGA broken, waiting for replacement.
15 Dec 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | Some smaller files on 10.12., might be due to snow. LGR pressure at 45.3 Torr
14 Dec 2021 | CH-CHA | MS | 0 | 0 | 0 |  | 0 | 0
13 Dec 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 2 | data gap since 8 Dec with lots of small files since then; at the moment system runs fine, most recent data are produced and stored
13 Dec 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
13 Dec 2021 | CH-DAV | PM | 0 | 0 | 1 |  | 0 | 0 | irgashow shows 93 % AGC.
10 Dec 2021 | CH-AWS | GA | 0 | 0 | 0 |  | 0 | 0
08 Dec 2021 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | Irga window dirtiness at 93.8%, probably raining. LGR pressure at 44.5 Torr
06 Dec 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 1 | small gap on Dec 1st and 2nd, likely due to bad weather or IRGA calibration, but system running fine at the moment
06 Dec 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | Signal strength dropped from 100% to currently 93.4% (still fine)
06 Dec 2021 | CH-DAS | LK | 0 | 0 | 0 |  | 0 | 0
02 Dec 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | LGR pressure at 44.45 Torr, .C00 files are smaller 20 MB instead of 26 MB, irga was removed for calibration but seems back now (data are incoming)
29 Nov 2021 | CH-DAV | MG | 0 | 0 | 1 |  | 0 | 0 | New computer running fine. IRGA flow rate lower than normal (13.2-13.5) but still within range. EC data gap between 8-19.11.2021 (due to IRGA and tower box issues - see previous notes and GIN)
29 Nov 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
28 Nov 2021 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0
28 Nov 2021 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
26 Nov 2021 | CH-AWS | GA | 0 | 0 | 1 |  | 0 | 0 | CO2(MFd) around 450
25 Nov 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0 | AGC 100%, thanks TB.
24 Nov 2021 | CH-CHA | IF | 0 | 0 | 2 |  | 0 | 0 | Negative concentrations incoming from LGR!
23 Nov 2021 | CH-DAS | LK | 0 | 0 | 0 |  | 0 | 0
22 Nov 2021 | CH-CHA | MS | 0 | 0 | 0 |  | 0 | 0
19 Nov 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
19 Nov 2021 | CH-AWS | GA | 0 | 0 | 0 |  | 0 | 0
17 Nov 2021 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | LGR pressure at 37.9 Torr
15 Nov 2021 | CH-DAS | LK | 0 | 0 | 0 |  | 0 | 0
15 Nov 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 2 | IRGA data coming in again, fixed by MS
15 Nov 2021 | CH-DAV | MG | 0 | 2 | 2 |  | 0 | 0 | Several issues that affected CH-DAV data since 8 November: 1) Power cut on the tower during the box refurbishment (affected EC and meteo-profile), 2) replaced IRGA (for yearly calibration) was factory calibrated and had latest firmware installed that is not recommended (currently being downgraded - IRGA data should be ignored since 10.11 until further notice) 3) LANNER failed again (probably faulty RAM banks) and needs to be replaced (work ongoing).
14 Nov 2021 | CH-FRU | LH | 0 | 2 | 2 |  | 2 | 2 | IRGA does not send data
14 Nov 2021 | CH-FRU | MS | 0 | 0 | 0 |  | 0 | 0 | irga is online again
14 Nov 2021 | CH-CHA | MS | 0 | 0 | 0 |  | 0 | 0
12 Nov 2021 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0 | Window little dirty. Will clean it next week
12 Nov 2021 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0 | All good.
12 Nov 2021 | CH-AWS | GA | 0 | 0 | 0 |  | 0 | 0
11 Nov 2021 | CH-OE2 | TB | 0 | 0 | 1 |  | 0 | 0 | AGC is at 80%
11 Nov 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
10 Nov 2021 | CH-DAV | MG | 0 | 0 | 2 |  | 1 | 0 | Issues with the EC after refurbishment of the tower box. Sonic running again. IRGA was replaced. The newly mounted IRGA (just arrived from LI-COR) has a firmware upgrade to LI-7200**RS** but this new firmware is not recommended by ETC and right now we are trying to downgrade it.
09 Nov 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | LGR pressure at 42.1 Torr after filter change
08 Nov 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
05 Nov 2021 | CH-DAS | LK | 0 | 0 | 0 |  | 0 | 0
05 Nov 2021 | CH-AWS | GA | 0 | 0 | 0 |  | 0 | 0
04 Nov 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
02 Nov 2021 | CH-OE2 | TB | 0 | 0 | 1 |  | 0 | 0 | AGC 80%
01 Nov 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
01 Nov 2021 | CH-DAV | PM | 0 | 0 | 0 |  | 0 | 0
28 Oct 2021 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | LGR pressure at 39.2 Torr
28 Oct 2021 | CH-DAS | LK | 0 | 0 | 0 |  | 0 | 0
25 Oct 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
25 Oct 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
25 Oct 2021 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | LGR pressure at 41.6. Change filter soon
21 Oct 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
21 Oct 2021 | CH-LAS | TB | 0 | 0 | 0 |  | 0 | 0
21 Oct 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
21 Oct 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
21 Oct 2021 | CH-DAS | LK | 0 | 0 | 0 |  | 0 | 0
21 Oct 2021 | CH-AWS | GA | 0 | 0 | 0 |  | 0 | 0
18 Oct 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
16 Oct 2021 | CH-AWS | GA | 0 | 0 | 0 |  | 0 | 0
14 Oct 2021 | CH-OE2 | RM | 0 | 0 | 1 |  | 0 | 0 | AGC 80%
14 Oct 2021 | CH-DAS | LK | 0 | 0 | 0 |  | 0 | 0
14 Oct 2021 | CH-OE2 | TB | 0 | 0 | 1 |  | 0 | 0 | AGC ~80%
14 Oct 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
14 Oct 2021 | CH-LAS | TB | 1 | 1 | 1 |  | 1 | 1 | new site computer
13 Oct 2021 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | LGR pressure at 42 Torr (usually 45 Torr). Need filter change
13 Oct 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
11 Oct 2021 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0
11 Oct 2021 | CH-AWS | LH | 0 | 0 | 0 |  | 0 | 0
11 Oct 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
07 Oct 2021 | CH-OE2 | RM | 0 | 0 | 1 |  | 0 | 0 | AGC 80%
07 Oct 2021 | CH-DAS | LK | 0 | 0 | 0 |  | 0 | 0
05 Oct 2021 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | Irga window slightly dirty ~81%, morning fog?
05 Oct 2021 | CH-FRU | MS | 0 | 0 | 0 |  | 0 | 0
04 Oct 2021 | CH-DAV | PM | 0 | 0 | 0 |  | 0 | 0
04 Oct 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
27 Sep 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
24 Sep 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | New file created on 23.9., so I checked the data. CH4 conc are quite high (3.1 ppm around 10:20) but else everything is working
22 Sep 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0 | The AGC and WD of both eddy system is back at a normal factor. Other than last week (was planning to fix it tomorrow)
22 Sep 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
22 Sep 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0 | AGC 86.7%
20 Sep 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
20 Sep 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | Power interruption, so some smaller files on 15./16.9.
20 Sep 2021 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0
20 Sep 2021 | CH-DAV | PM | 0 | 0 | 0 |  | 0 | 0
19 Sep 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
17 Sep 2021 | CH-LAE | TB | 0 | 0 | 1 |  | 0 | 0 | The AGC and the WD (Subcanopy) are not ideal. Subcanopy WD about 64% probably due to rainfall. AGC of T1 System about ~82 something. Planning to clean when problem persist
17 Sep 2021 | CH-OE2 | TB | 0 | 0 | 1 |  | 0 | 0 | AGC is low ~80
13 Sep 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0 | AGC 86.7%
09 Sep 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
08 Sep 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
08 Sep 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
07 Sep 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
06 Sep 2021 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0
03 Sep 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0 | AGC 93.4%
02 Sep 2021 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0 | All is good
02 Sep 2021 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 1 | The harddrive is 92% full. Needs to delete some data
01 Sep 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
01 Sep 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
31 Aug 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
25 Aug 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
25 Aug 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
23 Aug 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
23 Aug 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
17 Aug 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
16 Aug 2021 | CH-AWS | PM | 0 | 0 | 1 |  | 0 | 0 | IRGA without flow
16 Aug 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0 | some smaller files the last days
13 Aug 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | LGR cell pressure at 45 Torr
13 Aug 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0 | AGC 93.4%
12 Aug 2021 | CH-DAV | TB | 0 | 0 | 0 |  | 0 | 0 | IRGA slightly dirty and needs cleaning by the next fieldday
09 Aug 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
06 Aug 2021 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0 | All is good
06 Aug 2021 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0 | All is good
29 Jul 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
28 Jul 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 1 | Data gap on 25 and 26.7 because disk was full and files could not be written.
28 Jul 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
26 Jul 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | Change from Aerodyne QCL to LGR QCL last week. Data are incoming!
21 Jul 2021 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0 | All is good.
21 Jul 2021 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0 | Everything is fine.
19 Jul 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
19 Jul 2021 | CH-AWS | PM | 0 | 0 | 1 |  | 1 | 0 | irga data acquisition stopped on July 5. Program restart resolved issue
15 Jul 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | IRGA signal strength down to 86.7%.
13 Jul 2021 | CH-CHA | IF | 0 | 1 | 1 |  | 0 | 0 | high irga window dirtiness probably because of rain. sonicshow shows flag 40 for QCL --> means "Status ok, old data used".
12 Jul 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
12 Jul 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
12 Jul 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Last INO check :-(
01 Jul 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | Started a new flux file on 29.6. 14:00. QCL values ok, 2.8 V inlet valve, 855 mV signal, pressure 39.9 Torr. Files transferred again.
29 Jun 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
29 Jun 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Cell pressure 43.4 Torr
28 Jun 2021 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0
28 Jun 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
24 Jun 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | Data from station computer not transferred since 22.6. QCL inlet valve at 2.8 V, signal at 811 mV
24 Jun 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
24 Jun 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
21 Jun 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
21 Jun 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0 | AGC 93.4%
21 Jun 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Cell pressure 43.9 Torr
18 Jun 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
17 Jun 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
14 Jun 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
10 Jun 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | Looks good. QCL signal at 885 mV, inlet valve at 2.1 V
07 Jun 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
07 Jun 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
04 Jun 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 2 | 2 | Setup changed on 1.6., works so far but data transfer not. WE and TB are working on that. IRGA: flow a bit high (14L instead of 11-12L).
04 Jun 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Cell pressure 45 Torr (filters changed on 26.5.)
03 Jun 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
02 Jun 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL signal at ~900 mV, inlet valve: 2.36 V, pressure: 39.9 Torr
01 Jun 2021 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0 | IRGA 7200 running great now - Courtesy of Thomas and Phillip
01 Jun 2021 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0
28 May 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | Sonic failed on 19.05 and was restarted on 24.05. Cleaned the IRGA today and it is running fine now.
26 May 2021 | CH-LAE | TB | 0 | 0 | 2 |  | 0 | 0 | WTF. LI7200 is again on zero flow.
26 May 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
25 May 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
25 May 2021 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0
19 May 2021 | CH-LAE | TB | 0 | 0 | 2 |  | 0 | 0 | flow is too low. 3-4
19 May 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
19 May 2021 | CH-LAE | AS | 0 | 0 | 1 |  | 0 | 0 | The flowrate is very less, I think most probably due to rain.
19 May 2021 | CH-LAS | AS | 0 | 0 | 1 |  | 0 | 0 | The IRGA is a little dirty - cleaning it tomorrow
18 May 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | Signal strength 86.7%
18 May 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL filter changed 17.5.2021, 1.6 V inlet valve
14 May 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0 | AGC 86.7%
14 May 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Cell pressure 29.8 Torr. Change filter soon.
12 May 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
10 May 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0 | short data gap on 2 May during the night
06 May 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0 | Subcanopy IRGA: slightly dirty
06 May 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
04 May 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
04 May 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
03 May 2021 | CH-DAV | MG | 0 | 0 | 1 |  | 0 | 0 | IRGA signal strength has dropped from strong (93%) to good (86%).
30 Apr 2021 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | Window dirtiness at 93.8%, probably raining
28 Apr 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
28 Apr 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
27 Apr 2021 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0 | root partition was 100% full, but not /home. Some service dumped messages in the syslog
26 Apr 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
23 Apr 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0 | AGC 93.4%
23 Apr 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Cell pressure 39 Torr
23 Apr 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
21 Apr 2021 | CH-LAE | AS | 0 | 0 | 1 |  | 0 | 0 | AGC still 86%, but good signal.
21 Apr 2021 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0
18 Apr 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
16 Apr 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0 | AGC 93.4%
16 Apr 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Cell pressure 38.5 Torr
14 Apr 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
14 Apr 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
13 Apr 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
12 Apr 2021 | CH-FRU | LH | 0 | 0 | 1 |  | 0 | 0 | AGC: 93.8%
09 Apr 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0 | AGC 93.4%
09 Apr 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Cell pressure 39.8 Torr
08 Apr 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL: 2.67 iValve, 40 Torr, 1024 mV
06 Apr 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
01 Apr 2021 | CH-LAE | TB | 0 | 0 | 1 |  | 0 | 0 | AGC is 86%. Keep watching
01 Apr 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
29 Mar 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL running smoothly: 2.6 V ivalve, 39.9 Torr, signal 975 mV
26 Mar 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0 | AGC 93.4%
26 Mar 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Cell pressure 41 Torr
22 Mar 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0 | data are complete for the last weeks, seems like only the modem did not work during the previous check but data logging looks unaffected
22 Mar 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
19 Mar 2021 | CH-FRU | LH | 2 | [5 empty] | ssh connection to site did not work, but raw data on the server looks complete and was transferred normally last night, newest file is 2021031813.X00
19 Mar 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
19 Mar 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Cell pressure at 36 Torr
18 Mar 2021 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0 | All Good :)
18 Mar 2021 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0
17 Mar 2021 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | Irga window dirtiness at 93.6% (raining?). QCL inlet valve 2.1 V, 40 Torr. QCL signal below 1000 mV, clean mirrors?
16 Mar 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
16 Mar 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
12 Mar 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
12 Mar 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
08 Mar 2021 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0
05 Mar 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
05 Mar 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
04 Mar 2021 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0 | All Good :)
04 Mar 2021 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0 | All Good :)
03 Mar 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
03 Mar 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
03 Mar 2021 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | QCL pressure at 31 Torr (Sahara dust, pollen?), needs filter change
02 Mar 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
01 Mar 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | Sonic was replaced on 26.2.
26 Feb 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
26 Feb 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
24 Feb 2021 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | Irga window dirtiness at 93% at 09:00, might be fog. QCL inlet valve at 2.9 V
24 Feb 2021 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0 | Window dirtiness of IRGA was 93% in the morning, might be fog.
24 Feb 2021 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0 | Alles Gut!
23 Feb 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0 | happy.gif
19 Feb 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
19 Feb 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Cell pressure at 13 Torr
18 Feb 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
16 Feb 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
16 Feb 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
15 Feb 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | Cell pressure at 35 Torr, inlet valve at 5 V, needs filter change
15 Feb 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
12 Feb 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
12 Feb 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Cell pressure at 41 Torr
10 Feb 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
10 Feb 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
09 Feb 2021 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
09 Feb 2021 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0
08 Feb 2021 | CH-AWS | PM | 0 | 1 | 0 |  | 0 | 0 | sonic data corrupted due to snow-fall. otherwise running smoothly
08 Feb 2021 | CH-DAV | PM | 0 | 0 | 0 |  | 1 | 0 | binary EC files smaller due to lack of QCL data
08 Feb 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
04 Feb 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
04 Feb 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL inlet valve at 5V --> needs filter change. Irga concentrations and window dirtiness back to normal
03 Feb 2021 | CH-CHA | MS | 0 | 0 | 0 |  | 0 | 0
03 Feb 2021 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | (morning) negative H2O concentrations, window dirtiness at 68%
01 Feb 2021 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0
01 Feb 2021 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
29 Jan 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | IRGA signal strength dropped from 100% to 93.4%
28 Jan 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
28 Jan 2021 | CH-INO | RM | 0 | 0 | 1 |  | 0 | 0 | Cell pressure at 11 Torr
26 Jan 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL inlet valve at 0.45 V
26 Jan 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
26 Jan 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
25 Jan 2021 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0
25 Jan 2021 | CH-DAV | PM | 0 | 0 | 0 |  | 1 | 0 | QCL still not running. Binary filesize reduced.
25 Jan 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 1 | system running normally at the moment, but several interruptions over the last days (small files), likely due to weather
24 Jan 2021 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
24 Jan 2021 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0
22 Jan 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
20 Jan 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 2 | Data gap between 14-19 Jan, most likely due to snow. System running normally since 19 Jan 11:09h
19 Jan 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
19 Jan 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
18 Jan 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | Sonic did not work due to snow, ~3 days of data gap. QCL inlet valve at 3.22 V, 900 MB Memory
13 Jan 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
12 Jan 2021 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
12 Jan 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
11 Jan 2021 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0
11 Jan 2021 | CH-DAV | PM | 0 | 0 | 0 |  | 1 | 0 | QCL not running since 2020-12-29, thus EC filesize is reduced
11 Jan 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
11 Jan 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Cell pressure at 41 Torr. Files from 08.01.2021 in a separate folder (10_R3_LI75) but data complete for the last week.
08 Jan 2021 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
06 Jan 2021 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL inlet valve at 2.1 V, 40 Torr pressure, memory at 800 MB
04 Jan 2021 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
04 Jan 2021 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
04 Jan 2021 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 1 | Several smaller files on 01.01.2021 but data seems complete since then.
04 Jan 2021 | CH-LAE | TB | 0 | 0 | 2 |  | 0 | 0 | Irga was not running. Troubleshooting. Restartet computer, running now.
04 Jan 2021 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
29 Dec 2020 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | Could not connect to logger directly from home office, but fluxes look good for the past several weeks and data files are available. There are some gaps in the first half of December, since then files are complete
28 Dec 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
28 Dec 2020 | CH-INO | RM | 0 | 0 | 1 |  | 0 | 0 | Window dirty, might be due to snow. Cell pressure at 40 Torr.
22 Dec 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
17 Dec 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
16 Dec 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL inlet valve at 2.5 V after Andreas adjusted it. Deleted files on station computer.
15 Dec 2020 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0 | Mast was stuck (frozen) on 2020-12-14 from 8:00 to 13:30. Re-gained operation again independently at 14:00
15 Dec 2020 | CH-DAV | PM | 0 | 0 | 0 |  | 0 | 0
14 Dec 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
11 Dec 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
11 Dec 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
11 Dec 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL inlet valve at 3.2 V, pressure at 40 Torr, adjust next time
09 Dec 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
09 Dec 2020 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
07 Dec 2020 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0 | power outage on 2020-12-06. 1 m of snow, mast adapted correctly
07 Dec 2020 | CH-DAV | PM | 0 | 0 | 0 |  | 0 | 0
04 Dec 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
04 Dec 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 1 | No data transfer today. Check on Monday again. Cell pressure 37.5 Torr. Flux QVs not working since 23.11.
01 Dec 2020 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 1 | Window dirtiness irga 93%, probably because it is snowing. Files are a bit smaller for the last days due to irga calib and QCL work
01 Dec 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0 | On Friday the Flow was low (maybe it was foggy). Now everything seems normal
01 Dec 2020 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 1 | The harddisk usage is now 90%.
01 Dec 2020 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | can currently not connect to logger directly, but calculated fluxes look good and complete
30 Nov 2020 | CH-AWS | PM | 0 | 0 | 2 |  | 1 | 0 | IRGA was not delivering data since Nov 25. Restarting irga service resolved problem.
30 Nov 2020 | CH-DAV | PM | 0 | 0 | 0 |  | 0 | 1 | Lanner was 100% full. Moved data until 20201031 from OK to WIN PC
27 Nov 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
27 Nov 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Cell pressure at 40 Torr
24 Nov 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
24 Nov 2020 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0
23 Nov 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
23 Nov 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
20 Nov 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0 | IRGA calibrated this week. Thanks TB!
20 Nov 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
20 Nov 2020 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0 | IRGA exchanged on Tue 2020-11-17
20 Nov 2020 | CH-DAV | PM | 0 | 0 | 0 |  | 0 | 0
18 Nov 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
18 Nov 2020 | CH-OE2 | TB | 0 | 2 | 2 |  | 2 | 0 | IRGA calibration
17 Nov 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
13 Nov 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | Factory-calibrated IRGA running since 5.11.2020.
13 Nov 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
13 Nov 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
09 Nov 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
09 Nov 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
09 Nov 2020 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0
06 Nov 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
06 Nov 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | Good news: QCL is working again after PM installed a new driver. Lets see for how long.
03 Nov 2020 | CH-CHA | IF | 0 | 0 | 2 |  | 0 | 2 | No QCL data since 1st October.
28 Oct 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
28 Oct 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
27 Oct 2020 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
27 Oct 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
23 Oct 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
23 Oct 2020 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
22 Oct 2020 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0
22 Oct 2020 | CH-DAV | PM | 0 | 0 | 0 |  | 0 | 0
22 Oct 2020 | CH-CHA | MS | 0 | 1 | 0 |  | 1 | 0 | for sonicshow no qcl data. nor filesize data leak since 01 Okt. qcl is not running.
20 Oct 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 1 | Small data gap of approx 30 hours after 13 Oct
19 Oct 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
16 Oct 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | See GIN entry. QCL is working again since yesterday. QCL data gap from 1.10. onwards
15 Oct 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0 | fixed flow modul, see fieldbook. flow was ok the whole time
15 Oct 2020 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0 | fixed the sonic
14 Oct 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0 | The hard drive is now almost 90% full
14 Oct 2020 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0
13 Oct 2020 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0
12 Oct 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 2 | No or only few QCL data until 12th October, problems with laser switching off. Mike Moore from Aerodyne made the system run today, let's see for how long.
09 Oct 2020 | CH-DAV | PM | 0 | 0 | 0 |  | 0 | 0
09 Oct 2020 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0
08 Oct 2020 | CH-OE2 | RM | 0 | 2 | 0 |  | 0 | 0 | Sonic is again not working
08 Oct 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
08 Oct 2020 | CH-OE2 | TB | 0 | 2 | 0 |  | 0 | 0 | sonic has broken transducer pair one (for longer now) no sonic measurement
07 Oct 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | On 25 September sonic stopped, and needed a hard power reset.
05 Oct 2020 | CH-CHA | IF | 2 | 2 | 2 |  | 2 | 2 | Connection to station lost, no data transfer since 1.10.2020
28 Sep 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
25 Sep 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | Sonic was replaced on 22.09 for factory calibration.
25 Sep 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0 | All Good
25 Sep 2020 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0 | Irga needs cleaning. I will clean it next week on my visit to LAE
24 Sep 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL inlet valve at 2.8 V, Memory at 1 GB
24 Sep 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
24 Sep 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
22 Sep 2020 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
22 Sep 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
18 Sep 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
18 Sep 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
18 Sep 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
16 Sep 2020 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
16 Sep 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
16 Sep 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
13 Sep 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
10 Sep 2020 | CH-OE2 | RM | 0 | 2 | 0 |  | 0 | 1 | Again Sonic problems. Needs to be replaced.
10 Sep 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
08 Sep 2020 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | Irga window needs cleaning, maybe due to fog (at ~08:00), check later
08 Sep 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
08 Sep 2020 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
08 Sep 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
07 Sep 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
03 Sep 2020 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0 | no sonic problems occur
03 Sep 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
01 Sep 2020 | CH-OE2 | RM | 0 | 2 | 1 |  | 0 | 0 | Again Sonic problems: highly unrealistic values for u/v/w and temperature. IRGA needs cleaning.
01 Sep 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
31 Aug 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
31 Aug 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
30 Aug 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
28 Aug 2020 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | Irga at 62.5% window dirtiness, probably raining
25 Aug 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL inlet valve: 2V, memory: 1.15 GB
24 Aug 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
24 Aug 2020 | CH-OE2 | RM | 0 | 0 | 1 |  | 0 | 0 | AGC 86.7% -> IRGA needs cleaning
24 Aug 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
21 Aug 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL inlet valve at 1.4 V, Memory at 1.1 GB
17 Aug 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
17 Aug 2020 | CH-OE2 | RM | 0 | 0 | 1 |  | 0 | 0 | AGC 86.7% -> IRGA needs cleaning
17 Aug 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
17 Aug 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL Memory usage at 1.1 GB, inlet valve at 2.63 V
16 Aug 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
16 Aug 2020 | CH-LAS | AS | 0 | 0 | 0 |  | 0 | 0
12 Aug 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL: inlet valve at 2.34 V, laser current at 340 mA
11 Aug 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0 | logger has new IP address
10 Aug 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
10 Aug 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
10 Aug 2020 | CH-DAV | PM | 0 | 0 | 0 |  | 0 | 0
10 Aug 2020 | CH-AWS | PM | 0 | 0 | 1 |  | 0 | 0 | IRGA AGC on 86.7 %. Cleaning advised
08 Aug 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
05 Aug 2020 | CH-OE2 | MS | 0 | 0 | 1 |  | 0 | 0 | Irga needs to be cleaned
05 Aug 2020 | CH-INO | MS | 0 | 0 | 0 |  | 0 | 0
05 Aug 2020 | CH-CHA | MS | 0 | 0 | 0 |  | 0 | 0
03 Aug 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
03 Aug 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
03 Aug 2020 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | Irga window sligthly dirty (probably due to rain), QCL inlet valve at 2V, 40 Torr pressure, 1GB memory used.
28 Jul 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
27 Jul 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
27 Jul 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
27 Jul 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | Past files are smaller due to before-mentioned issues
26 Jul 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
24 Jul 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | After some QCL problems (external harddisk, pump) this week, the pump got exchanged on 23.7.2020 and the HDD connection fixed (see GIN), files are smaller. Now, everything is working again.
22 Jul 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
21 Jul 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
21 Jul 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
17 Jul 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
17 Jul 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
16 Jul 2020 | CH-CHA | MS | 0 | 0 | 0 |  | 0 | 1 | Due to lightning protection upgrade. PWD had to be shut down for a moment. Therefore visvais 202007111200.dat is smaller
16 Jul 2020 | CH-OE2 | MS | 0 | 0 | 0 |  | 0 | 0
16 Jul 2020 | CH-INO | MS | 0 | 0 | 0 |  | 0 | 0 | Deleted data on the LaserPC from 2018 to have more storage space again
15 Jul 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | Files smaller from 09.07. - 11.07.2020 because of TDLWintel error (external hard drive lost connection). QCL memory usage: 0.98 GB, CPU: 34%, cell pressure: 39.4 Torr, Laser current: 633 mA
13 Jul 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
08 Jul 2020 | CH-FRU | MS | 0 | 0 | 0 |  | 1 | 0 | Because of the temporary shutdown of the sites due to maintenance work in Früebüel there is a small loss of data
07 Jul 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
07 Jul 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
07 Jul 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
07 Jul 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
06 Jul 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
06 Jul 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
05 Jul 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
02 Jul 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | Past files are smaller due to pump shut down. Else everything fine again.
29 Jun 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
29 Jun 2020 | CH-CHA | PM | 0 | 1 | 0 |  | 1 | 0 | no QCL data
29 Jun 2020 | CH-AWS | PM | 0 | 0 | 0 |  | 0 | 0
29 Jun 2020 | CH-DAV | PM | 0 | 0 | 0 |  | 0 | 0
29 Jun 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 2 | On Friday evening, the QCL pump shut down. Thus no data over the weekend. Went there today and restarted QCL and pump. Fine again.
25 Jun 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
24 Jun 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | Laser current around 600 mA
24 Jun 2020 | CH-OE2 | RM | 0 | 2 | 0 |  | 0 | 1 | u, v, w and temperature are jumping and values are highly unrealistic. Moreover, many small files instead of 4 complete files over the day.
24 Jun 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
23 Jun 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 1 | 5th June 23-6th June 7:00, 7th June 14:30- 8th June 1:00 we have a data loss
23 Jun 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
20 Jun 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
15 Jun 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
15 Jun 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
12 Jun 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
12 Jun 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
11 Jun 2020 | CH-OE2 | TB | 0 | 0 | 0 |  | 1 | 0 | licor0.log was too big. Needed to restart
11 Jun 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
10 Jun 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | LI7200 replaced today with a spare lab-calibrated LI7200 to be sent for repair.
05 Jun 2020 | CH.LAE | AS | 0 | 0 | 0 |  | 0 | 0 | All Good
05 Jun 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | On 04.06.2020 files are smaller. Probably due to MS station visit. QCL current still at 940 mA, signal is 1970 mV
04 Jun 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 1 | May 20 19:00 to May 21 13:00, we have data gap
03 Jun 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
03 Jun 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | running out of disk space (should be still enough for 2-3 weeks)
03 Jun 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
03 Jun 2020 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
03 Jun 2020 | CH_LAE | TB | 0 | 0 | 0 |  | 0 | 0
02 Jun 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL Laser current at 970 mA
28 May 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | Deleted files on the CHA computer to free disk space, usage of home now: 53%. QCL memory usage at 890 MB after restart on 19.05.2020
27 May 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 1 | The EC data from yesterday (since 11:00) have columns shifted. Now files are being logged correctly again.
25 May 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
25 May 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | cell pressure at 49 Torr
25 May 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
24 May 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0 | All good here
23 May 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 1
19 May 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
18 May 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | Sligthly changed QCL data acquisition (see GIN). Everything looks good. QCL Memory at 1.1 GB
15 May 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0 | All good
15 May 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
15 May 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Moved INO from OE2 to AES on 12.5.2020. Everything works and looks good.
12 May 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
12 May 2020 | CH-OE2 | TB | 0 | 0 | 1 |  | 0 | 0 | nothing new
12 May 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0 | 29th April and 1st May loss of data
11 May 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
11 May 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
08 May 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
07 May 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
06 May 2020 | CH-OE2 | RM | 0 | 0 | 1 |  | 0 | 0 | Over the last weeks, I realized that during rain AGC of the IRGA drops to 0% because of water inside but we still don't know why. Besides that, data looks good and is complete.
06 May 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
05 May 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 1 | 0 | between 30th Apr and 1st Mai, the computer time was shifted and some files have more minutes. the affected files should be removed from calculations. an Enrty added in GIN by PM.
05 May 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0 | would be good to delete old OK folder (2018)
05 May 2020 | CH-OE2 | TB | 0 | 0 | 1 |  | 0 | 0 | when it is raining the AGC is not ok. There is clearly a issue with the LI7200 which we need to address
04 May 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
29 Apr 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
29 Apr 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0 | All good
27 Apr 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
24 Apr 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
24 Apr 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
21 Apr 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | After QCL filter change cell pressure is at 40 Torr again. Memory usage at 840 MB.
21 Apr 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
21 Apr 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
21 Apr 2020 | CH-OE2 | TB | 0 | 0 | 1 |  | 0 | 0 | LI7200 with broken outlet thermocouple since 18.04.2020
20 Apr 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
20 Apr 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
20 Apr 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
20 Apr 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
14 Apr 2020 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | QCL cell pressure at 18 Torr, change filter?
14 Apr 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
13 Apr 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
08 Apr 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | Everything is fine. Checked frequency of QCL in binaries ~10 Hz
07 Apr 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
06 Apr 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
06 Apr 2020 | CH-LAE | TB | 0 | 0 | 1 |  | 0 | 0 | signal is low. irga needs cleaning. difficult at the moment
06 Apr 2020 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
03 Apr 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
01 Apr 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
31 Mar 2020 | CH-CHA | IF | 0 | 1 | 2 |  | 0 | 2 | QCL pump shut down on 30.3. 16:00. Therefore no QCL data. Philip went to restart it. Now it is working again. But with sonicshow every ~15th record says NO QCL DATA --> check QCL frequency in binary files if transferred.
31 Mar 2020 | CH-LAE | TB | 0 | 0 | 1 |  | 0 | 0 | AGC is fluctuating. needs additional attention
30 Mar 2020 | CH-AWS | SG | 0 | 0 | 1 |  | 0 | 0
27 Mar 2020 | CH-LAE | AS | 0 | 0 | 0 |  | 0 | 0
27 Mar 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
27 Mar 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
27 Mar 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
23 Mar 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
20 Mar 2020 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | Very high CH4 (3700) and N2O (338) conc. On 19.3. they fertilized Parcel 4. Irga window dirtiness at 93.3%, maybe due to fog? Will check again later --> window dirtiness at 50% again, conc of CH4 and N2O lower --> boundary layer transition
20 Mar 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
20 Mar 2020 | CH-INO | RM | 0 | 0 | 1 |  | 0 | 0 | Cell pressure currently at 40 Torr
20 Mar 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
16 Mar 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
15 Mar 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
13 Mar 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
13 Mar 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
13 Mar 2020 | CH-INO | RM | 0 | 0 | 1 |  | 0 | 0 | Cell pressure is at 38 Torr
12 Mar 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
11 Mar 2020 | CH-DAV | MG | 0 | 2 | 0 |  | 2 | 2 | Had to re-power the sonic as it stopped working yesterday 18:00.
11 Mar 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
10 Mar 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
09 Mar 2020 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
04 Mar 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | QCL is on-site and running. However the memory ran full over the weekend. Thus files are smaller. Keep an eye on it.
03 Mar 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 1 | We had a gap (likely power cut) between 20:30 yesterday and today 11:00 AM. EC system running fine again.
02 Mar 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
02 Mar 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
02 Mar 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
28 Feb 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | Still no QCL data. We will install it today again.
25 Feb 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
25 Feb 2020 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
24 Feb 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
24 Feb 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 1 | sonic again running normally; in summary no data b/w 12 and 19 Feb
24 Feb 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
21 Feb 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
21 Feb 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
17 Feb 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0 | LAS online again
17 Feb 2020 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
17 Feb 2020 | CH-FRU | LH | 0 | 2 | 2 |  | 0 | 0 | no IRGA data since 12 Feb
14 Feb 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
12 Feb 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
12 Feb 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
10 Feb 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | See previous check. sonicshow says "NO QCL DATA" which is to be expected.
07 Feb 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
07 Feb 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | Files are smaller because QCL is currently at ETH.
06 Feb 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
04 Feb 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
04 Feb 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
03 Feb 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
03 Feb 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
31 Jan 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
30 Jan 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0 | LAS offline. AGC low but ok
30 Jan 2020 | CH-OE2 | TB | 0 | 0 | 1 |  | 0 | 0
30 Jan 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
28 Jan 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | Due to the TDLWintel restart, file 2020012407.C00 is slightly smaller.
28 Jan 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
24 Jan 2020 | CH-CHA | IF | 0 | 0 | 2 |  | 0 | 0 | qclshow showed negative concentrations for N2O and CH4. Restarting TDLWintel fixed the problem
21 Jan 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 1 | 10th of anuary we have a gap
21 Jan 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL: CPU usage ~35%, Memory ~1.1GB
21 Jan 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
20 Jan 2020 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
20 Jan 2020 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
20 Jan 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
20 Jan 2020 | CH-OE2 | TB | 0 | 0 | 1 |  | 0 | 0 | Inlet thermocouple broke
20 Jan 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0 | LAS not online. AGC is low
13 Jan 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
09 Jan 2020 | CH-OE2 | TB | 0 | 0 | 1 |  | 0 | 0 | The Inlet Thermocouple of the LI7200 is broken
09 Jan 2020 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 1 | Internet again. Halleluja. We need to set up the automatic data download again.
09 Jan 2020 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
06 Jan 2020 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 1 | 21st September we have files that are only 29 kb. so, no file for this day. what happened on the site?
06 Jan 2020 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | 2019123107. C00 and 2019123113.C00 are slightly smaller (normally files are around 21 MB) --> Markus was on-site
02 Jan 2020 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
02 Jan 2020 | CH-DAV | LH | 1 | 1 | 1 |  | 1 | 1 | connection via SSH did not work with credentials from GIN, but newest 2020 files are on our server
23 Dec 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
23 Dec 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
20 Dec 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
19 Dec 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
18 Dec 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | Deleted files that were older than 180 days in the OK.2019 folder. Current home usage: 60%. High CH4 (> 2000) and N2O (~334) concentrations. QCL memory: ~900 MB
18 Dec 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
18 Dec 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | cell pressure of the laser is at 40 Torr
18 Dec 2019 | CH-LAE | TB | 2 | 0 | 0 |  | 1 | 1 | no internet. Sonic shows strange restarts. Sonic will be replaced
17 Dec 2019 | CH-CHA | MS | 0 | 0 | 0 |  | 0 | 0
17 Dec 2019 | CH-FRU | MS | 0 | 0 | 0 |  | 0 | 0
13 Dec 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
13 Dec 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL Memory at 820 MB. Home usage of 85% on the station computer. Remember to delete some old data before Christmas.
13 Dec 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
10 Dec 2019 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0 | comment on temp fluct by RM. It's possible that I just missed that fluctuation but it was there.
10 Dec 2019 | CH-LAE | TB | 2 | 0 | 0 |  | 1 | 1 | no internet. Small files, more than usual due to bad weather
09 Dec 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
09 Dec 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
09 Dec 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
09 Dec 2019 | CH-OE2 | RM | 0 | 2 | 0 |  | 0 | 0 | Sonic temperature is fluctuating largely within four measurements (6-10°C).
09 Dec 2019 | CH-INO | RM | 0 | 2 | 0 |  | 0 | 0 | sonic temperature is fluctuating
07 Dec 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
05 Dec 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | QCL Pump switched off on 29.11. ~12:00 until around 16:00. Discard QCL measurements in this period
03 Dec 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0 | Data from 5th of Nov and 21st Nov should be ignored, Technicians at the site.
03 Dec 2019 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
03 Dec 2019 | CH-LAE | TB | 2 | 0 | 0 |  | 1 | 0 | no internet.Small files, more than usual due to bad weather
02 Dec 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
02 Dec 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
02 Dec 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
02 Dec 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
29 Nov 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
28 Nov 2019 | CH-CHA | IF | 0 | 0 | 1 |  | 1 | 1 | Irga gets calibrated so no Irga data. Therefore also files are smaller. QCL and sonic still measuring.
28 Nov 2019 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
25 Nov 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
22 Nov 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL Memory: 950 MB, usage seems to slowly increase.
22 Nov 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 1 | webcam still not working
22 Nov 2019 | CH-INO | RM | 0 | 0 | 1 |  | 0 | 0 | Cell pressure of the LGR is slightly decreasing (41 Torr)
21 Nov 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
20 Nov 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
20 Nov 2019 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
15 Nov 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
13 Nov 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
11 Nov 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 1 | webcam still not working
11 Nov 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
08 Nov 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 1 | 5th of November (8:12-13) techs working on the site, file size is small. 6th of November (11.46-19) 3 files with smaller than usual size.
07 Nov 2019 | CH-OE2 | TB | 0 | 0 | 0 |  | 0 | 0
06 Nov 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
05 Nov 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
05 Nov 2019 | CH-LAE | TB | 2 | 0 | 0 |  | 0 | 0 | The site is not online due to ISDN problems
04 Nov 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
01 Nov 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
01 Nov 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
31 Oct 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
31 Oct 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | Restarted the QCL on 30.10.2019 13:40 MEZ so the file is slightly smaller then.
30 Oct 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
29 Oct 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
24 Oct 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0 | Since 22.10. everything is working again except the webcam.
24 Oct 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
23 Oct 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL: CPU usage 34%, Memory 890 MB
23 Oct 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
21 Oct 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
21 Oct 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 1 | just one file on 16th October is smaller than usual
21 Oct 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
18 Oct 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | QCL: CPU usage 34%, Memory 888 MB
14 Oct 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
11 Oct 2019 | CH-OE2 | RM | 2 | 2 | 2 |  | 2 | 2 | No connection to the site and no data since 2.10.2019. MS and ISG informed and working on the problem.
11 Oct 2019 | CH-INO | RM | 0 | 0 | 1 |  | 0 | 0 | CO2 concentration is very high and the mirror needs to be cleaned.
08 Oct 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
07 Oct 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
07 Oct 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
04 Oct 2019 | CH-OE2 | RM | 2 | 2 | 2 |  | 2 | 2 | No connection to the site and no data since 2.10.2019. MS and ISG informed and working on the problem.
04 Oct 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
03 Oct 2019 | CH-AWS | TB | 2 | 0 | 0 |  | 0 | 0 | network/firewall problem. On site repair. Could check measurement. Locally ok.
03 Oct 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0 | Ask Werner to remove old files. Soon we will have more free memory
01 Oct 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
01 Oct 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
01 Oct 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
30 Sep 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
27 Sep 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 1 | Had to restart the IRGA due to incomplete and smaller datafiles since yesterday.
27 Sep 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | data set after power cut complete!
27 Sep 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
25 Sep 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
25 Sep 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
25 Sep 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | Networs issue recently, data files on-site are OK but transfer to ETH server and to ICOS interrupted.
23 Sep 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
22 Sep 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
20 Sep 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 1 | Power cut --> data gap of two days. Already fixed again.
20 Sep 2019 | CH-INO | RM | 2 | 2 | 2 |  | 2 | 2 | No remote connection to the router possible since the power cut on 15.9.2019.
19 Sep 2019 | CH.LAE | TB | 0 | 0 | 0 |  | 0 | 0
17 Sep 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
17 Sep 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
16 Sep 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
16 Sep 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
16 Sep 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 2 | 0 | on the 8th of September (06.59) the file sizes are only 29 till 8th September at 19 also we have some small files on 11th and 12th
13 Sep 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
12 Sep 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
11 Sep 2019 | CH-OE2 | RM | 0 | 0 | 1 |  | 0 | 0 | IRGA needs cleaning soon
11 Sep 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
10 Sep 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
10 Sep 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0 | The root partition was full. Autocompletion therefore wasn't working anymore. Deleted old kernel versions in /lib/modules/ that were not necessary anymore (and clogging up the drive).
09 Sep 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
29 Aug 2019 | CH-OE2 | MS | 0 | 0 | 0 |  | 0 | 0
29 Aug 2019 | CH-INO | MS | 0 | 0 | 0 |  | 0 | 0
28 Aug 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | I can check via SSH (grassland server) that EC devices are running OK and data is being logged. Data however is still stored only on the computer in Davos.
27 Aug 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
27 Aug 2019 | CH-DAV | MG | 2 | 2 | 2 |  | 2 | 2 | Since the router update by empa on 20.08.2019 we have no direct access to the data in Davos. Currently data is being stored on the computer in Davos (LH checked via teamviewer). PM is working to restore the connection then we can transfer the data to ETH.
26 Aug 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
26 Aug 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
26 Aug 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
26 Aug 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
23 Aug 2019 | CH-AWS | SG | 0 | 0 | 2 |  | 0 | 0 | Couldnt check the irgashow, says queueid 73
20 Aug 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
20 Aug 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
19 Aug 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
19 Aug 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
17 Aug 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
16 Aug 2019 | CH-AWS | SG | 0 | 0 | 1 |  | 0 | 1 | 8th of August file size is smaller than usual- signal strength is 100.1
15 Aug 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
14 Aug 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
13 Aug 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
12 Aug 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
12 Aug 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
09 Aug 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
07 Aug 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | QCL pump did not work intermittently (lightning?). Smaller files on 06.08.2019
05 Aug 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
05 Aug 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
05 Aug 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
05 Aug 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
02 Aug 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
31 Jul 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 1 | I restarted the IRGA because of incomplete datasets since yesterday.
31 Jul 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
30 Jul 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
30 Jul 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
30 Jul 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
26 Jul 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
26 Jul 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
26 Jul 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
22 Jul 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
22 Jul 2019 | Ch-FRU | PL | 0 | 0 | 0 |  | 0 | 0
19 Jul 2019 | CH-AWS | SG | 0 | 0 | 2 |  | 2 | 0 | I cannot check the irgashow, tells me doesnt run under Mac- on 16th July we have a data gap (file size is small) between 7-19
19 Jul 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
19 Jul 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
18 Jul 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
18 Jul 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
16 Jul 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
16 Jul 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
15 Jul 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
15 Jul 2019 | CH-CHA | IF | 0 | 2 | 0 |  | 1 | 0 | qclshow not working, sonicshow NO QCL DATA! When I was on-site TDLWintel showed a memory error. Restart fixed the problem. First file with smaller filesize: 2019071501.C00
12 Jul 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
12 Jul 2019 | CH-OE2 | RM | 0 | 1 | 0 |  | 0 | 1 | Restarted IRGA at 12:07 due to NO IRGA DATA. Data gap of one day.
12 Jul 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
10 Jul 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
09 Jul 2019 | CH-OE2 | MS | 0 | 0 | 0 |  | 0 | 0
09 Jul 2019 | CH-INO | MS | 0 | 0 | 1 |  | 0 | 0 | IrgaWindow at 93.4%
08 Jul 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
08 Jul 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
08 Jul 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
03 Jul 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
02 Jul 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
02 Jul 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
01 Jul 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | 2019063019.C00 and 2019070101.C00 are too small. Probably TDLwintel stopped again because of high temperatures.
28 Jun 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
28 Jun 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
28 Jun 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
28 Jun 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
28 Jun 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
24 Jun 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
21 Jun 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
21 Jun 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
21 Jun 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
21 Jun 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
20 Jun 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
17 Jun 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | Binary files have not been transferred to archive since 13th June. Talked to Werner and changed the script, was running for Oensingen only for test purposes. Other sites also affected by that, solved now. QCL Teamviewer not running since last week, data are incoming though.
16 Jun 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
14 Jun 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
12 Jun 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
12 Jun 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 1 | Uncomplete files from June6 to June 7
12 Jun 2019 | CH-INO | MS | 0 | 0 | 0 |  | 0 | 0
12 Jun 2019 | CH-OE2 | MS | 0 | 0 | 1 |  | 0 | 0 | IRGA AGC at 93.4%. Window should be cleaned.
12 Jun 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
12 Jun 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
11 Jun 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
10 Jun 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
07 Jun 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 1 | 1 | After SRP-visit and check of QCL from Aerodyne on June 6th, the TDLWintel settings were set to 1 Hz instead of 10 Hz temporarily and also restarts of TDLWintel and the QCL happened, which probably caused a lower file size of the binary files (17 MB instead of 21 MB).
06 Jun 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
03 Jun 2019 | CH-OE2 | MS | 0 | 0 | 0 |  | 0 | 0
03 Jun 2019 | CH-INO | MS | 0 | 0 | 0 |  | 0 | 0
31 May 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
31 May 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
31 May 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
29 May 2019 | CH-FRU | LH | 0 | 0 | 0 |  | 0 | 0
29 May 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
27 May 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
27 May 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
27 May 2019 | CH-OE2 | MS | 0 | 0 | 0 |  | 0 | 0
27 May 2019 | CH-INO | MS | 0 | 0 | 0 |  | 0 | 0
24 May 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 2 | Sonic cable got fixed. However there is a data gap for flux data from 19.05.2019 to 23.05.2019.
24 May 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
24 May 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
24 May 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
21 May 2019 | CH-CHA | PL | 0 | 2 | 0 |  | 2 | 2 | No sonic data arriving at lanner/(logger) since 19.05.19.
21 May 2019 | CH-FRU | PL | 0 | 0 | 1 |  | 0 | 0 | Rain / bad signal strenght
20 May 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
17 May 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
17 May 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
16 May 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
16 May 2019 | CH-AWS | TB | 0 | 0 | 0 |  | 0 | 0
15 May 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | Data acquisition runs normally. But calculating fluxes (from the binary files) does not work on the RDS.
15 May 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
15 May 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
14 May 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
14 May 2019 | CH-DAV | MG | 2 | 2 | 2 |  | 2 | 2 | On 20190510 13:00 data acquisition stopped (sonicreadHS was running but sonicshow wasnt showing anything) until the next day that PM restarted it at 20190511 13:30. The script (ICOS-check-EC) that restarts if logging stops did not pick this because there was an error in the script which PM has fixed now. Data acquisition is back to normal now.
13 May 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
12 May 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
12 May 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | Station was moved to Aeschi and is measuring there since 10.5.2019.
10 May 2019 | CH-AWS | TB | 0 | 0 | 0 |  | 0 | 0
10 May 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
08 May 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
08 May 2019 | CH-CHA | PL | 0 | 0 | 1 |  | 0 | 0 | Irga shows slightly dirty (probably due to rain)
07 May 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | 91% percent usage of home. OK.2018 needs to be deleted. Who can help with checking? QCL issue was solved on 30th of April
07 May 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
06 May 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
03 May 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
29 Apr 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
29 Apr 2019 | CH-CHA | PL | 0 | 1 | 0 |  | 0 | 0
29 Apr 2019 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | not checked on logger, but fluxes look OK and complete in EddyPro results
29 Apr 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
29 Apr 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
26 Apr 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
26 Apr 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
25 Apr 2019 | CH-CHA | IF | 0 | 1 | 0 |  | 0 | 0 | 90% usage of home, remember to delete OK.2018 soon (after a last check that all of the data has been transferred to the archive). sonicshow still shows NO QCL DATA intermittently. This is because the measurement frequency is 4.7 instead of 10 Hz. Aerodyne is informed and tries to solve the issue. TV is still down
25 Apr 2019 | CH-CHA | PL | 0 | 1 | 0 |  | 0 | 2 | No QCL data from 19 april - 23 april
25 Apr 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
24 Apr 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
22 Apr 2019 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | not checked on logger, but flux calcs look OK and complete since 10 Apr
17 Apr 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 1 | visvais files get transferred to archive but not to OK.2019 folder --> talked to Werner, should be fixed
16 Apr 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 1 | Between 13.4. and 16.4 filesize of EC data is smaller
15 Apr 2019 | CH-CHA | PL | 0 | 1 | 0 |  | 0 | 0
15 Apr 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
15 Apr 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
15 Apr 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
10 Apr 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 1 | Problem with sonic files on April 5 from 10:00 - 13:00. Multiple files with only small size
10 Apr 2019 | CH-CHA | PL | 0 | 1 | 0 |  | 0 | 0
10 Apr 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0 | nothing new
09 Apr 2019 | CH-CHA | IF | 0 | 1 | 1 |  | 0 | 2 | No QCL data every 10th sonicshow record, irgashow shows high window dirtiness, visvais files did not get transferred to archive since 27th February (not moved to OK. folder) --> talk to Werner
05 Apr 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
03 Apr 2019 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | not checked on logger, but fluxes look complete, although file sizes were reduced on 25-26 Mar
01 Apr 2019 | CH-CHA | PL | 0 | 1 | 0 |  | 0 | 0 | Sonic show has no QCL data every couple of lines
01 Apr 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 1 | Uncomplete files from 17.03.19 - 18.03.19 and 25.03.19-26.03.19
01 Apr 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
29 Mar 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
28 Mar 2019 | CH-CHA | IF | 0 | 1 | 0 |  | 0 | 0
28 Mar 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0 | LI7200 is reinstalled and data from irgashow seems to be fine
26 Mar 2019 | CH-FRU | PL | 0 | 0 | 2 |  | 1 | 1 | Irga was not able to start up by it self after the power grid maintainance. Running since 09:50
26 Mar 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
26 Mar 2019 | CH-AWS | SG | 2 | 2 | 2 |  | 2 | 2 | the computer is down. it is more or less the same time as it happened last year. can it be due to temperature?
25 Mar 2019 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | not checked on logger, but data are complete in prelim (level-0) fluxes
25 Mar 2019 | CH-FRU | PL | 2 | 2 | 2 |  | 2 | 2 | CH-FRU was offline from 09:30 - 15:35 due to power grid maintainance.
22 Mar 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
22 Mar 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
21 Mar 2019 | CH-CHA | IF | 0 | 1 | 0 |  | 0 | 0 | same as last check, no QCL data every 10th sonicshow record
19 Mar 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 2 | 2 | EC data acquisition stopped at 19.03.2019 22:00 and resumed at 20.03.2019 01:30. SonicreadHS stdout: unable to open output file 2019032001.X00 neither in the data path nor in the current directory. Could not open sonicavg.dat. EC data acquisition is now in order.
18 Mar 2019 | CH-LAE | EPL | 0 | 0 | 2 |  | 0 | 0 | IRGA data is problematic. The LI-7200 is ready to be reinstalled on Wednesday.
13 Mar 2019 | CH-OE2 | MS | 0 | 0 | 1 |  | 0 | 0 | No IRGA Data between 6 - 7 March 2019
13 Mar 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
11 Mar 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
11 Mar 2019 | CH-LAE | EPL | 0 | 0 | 1 |  | 0 | 0 | IRGA AGC too low
10 Mar 2019 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | not checked on logger, but EC fluxes from calculations look complete for the year so far
08 Mar 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
08 Mar 2019 | CH-CHA | IF | 0 | 1 | 0 |  | 0 | 0 | same as in previous check, no QCL data every ~10th record of sonicshow. --> ask Lukas/Werner if normal (I think so because the QCL is recording in 10 Hz, the sonic in 20 Hz)
06 Mar 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
06 Mar 2019 | CH-CHA | PL | 0 | 1 | 0 |  | 0 | 1 | No QCL data every ~10 entries. Small files on feb 28
04 Mar 2019 | CH-LAE | EPL | 0 | 0 | 1 |  | 0 | 0 | IRGA AGC too low
04 Mar 2019 | CH-OE2 | MS | 0 | 0 | 1 |  | 1 | 0 | No irga data since 03.03.19 means data loss in packages 2019030301 - 2019030407
02 Mar 2019 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | not checked on logger, but EC fluxes from calculations are complete for the year so far
01 Mar 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
28 Feb 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
26 Feb 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | No automatic fluxes for CO2 for the recent days, QCL fluxes working
26 Feb 2019 | CH-INO | MS | 0 | 2 | 2 |  | 2 | 2 | Complete ShutDown of the mobile EddyStation due to 1.Rotation for InnoFarm project.
26 Feb 2019 | CH-OE2 | MS | 0 | 0 | 0 |  | 1 | 0 | Accidently powershutdown due to Maintenance of Site. So some data are missing
25 Feb 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
25 Feb 2019 | CH-LAE | EPL | 0 | 0 | 1 |  | 0 | 0 | IRGA AGC too low
22 Feb 2019 | CH-INO | MS | 0 | 0 | 0 |  | 0 | 0
22 Feb 2019 | CH-OE2 | MS | 0 | 0 | 1 |  | 1 | 0 | no irga data between 17 - 22 Feb. 2019
20 Feb 2019 | CH-AWS | IB | 0 | 0 | 0 |  | 0 | 0 | Snowheight.dat file OK.
18 Feb 2019 | CH-CHA | PL | 0 | 2 | 0 |  | 2 | 2 | No QCL Data incoming
18 Feb 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
18 Feb 2019 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | from auto fluxes, EC data look complete and OK
15 Feb 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
14 Feb 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
12 Feb 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0 | file size on 06.02.2019 between 17.21 to 19 is small
12 Feb 2019 | CH-CHA | IF | 0 | 2 | 0 |  | 2 | 2 | NO QCL DATA! Restarted qclread on logger, no effect. Must have something to do with the QCL computer. Since 11.02.2019, binary files are too small (no QCL data)
12 Feb 2019 | CH-INO | MS | 0 | 0 | 0 |  | 0 | 0
12 Feb 2019 | CH-OE2 | MS | 0 | 0 | 0 |  | 0 | 0
11 Feb 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
11 Feb 2019 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0 | IRGA from profile is not working.
08 Feb 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
08 Feb 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | CO2 and H2O fluxes are very noisy over the last two weeks
07 Feb 2019 | CH-INO | MS | 0 | 0 | 0 |  | 0 | 0
06 Feb 2019 | CH-OE2 | MS | 0 | 0 | 0 |  | 0 | 0
05 Feb 2019 | CH-LAE | EPL | 0 | 0 | 1 |  | 0 | 0 | IRGA AGC too low
03 Feb 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 2 | 2 | EC data acquisition stopped at 02.02.2019 03:00 and automatically resumed from 02.02.2019 11:00. Data acquisition and transfer is in order now.
01 Feb 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
01 Feb 2019 | CH-INO | RM | 0 | 0 | 1 |  | 0 | 0 | Negative H2O concentrations, other values are ok.
31 Jan 2019 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | no direct check on logger, but EC data look complete in auto-flux QVs
31 Jan 2019 | CH-LAE | TB | 0 | 0 | 1 |  | 0 | 0 | At the moment the AGC is to high. Snowfall. Checked with EPL days before and there it was ok.
31 Jan 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
30 Jan 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 2 | 2 | EC data were not logged between approx. 22:30 on 29.01.2019 and 1:30 on 30.01.2019. Seems like the script stopped and then restarted after 3 hours. Data acquisition and transfer in order now.
30 Jan 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
28 Jan 2019 | CH-LAE | EPL | 0 | 0 | 2 |  | 0 | 0 | Replacement IRGA seems defective. Negative H20 concentrations... Numbers are very different than from CH-LAS running at the same site.
28 Jan 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
28 Jan 2019 | CH-INO | RM | 0 | 0 | 2 |  | 0 | 0 | Negative/unrealistic CO2 concentrations and no values for the other variables. Restarted IRGA, CO2 still strange but this could be due to rain/snow (?)
23 Jan 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
22 Jan 2019 | CH-AWS | IB | 0 | 0 | 0 |  | 0 | 0 | The snowheight data (/home/data/CH-AWS-iDAQ-BOX1-0-1/CH-AWS_snowheight_201901.dat) was missing since 17.01.2019. The reason was a power interruption on that day and the Snowheight sensor did not reboot. Fixed the problem (IoLogik preset wrong for SHM50).
22 Jan 2019 | CH-CHA | IF | 0 | 2 | 1 |  | 1 | 0 | sonicshow: NO QCL DATA!! checked binary files and apparently this was already the case for 2019012013.C00. SInce then files are too small. Irgashow : Window dirtiness 43.8% TOO LOW
21 Jan 2019 | CH-LAE | EPL | 0 | 0 | 2 |  | 0 | 0 | IRGA AGC TOO LOW (43.8%)
20 Jan 2019 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | no direct check on logger, but fluxes look good and complete until 19 Jan
18 Jan 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 1 | Reduced filesize inbetween Jan 6 - Jan 8
18 Jan 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
18 Jan 2019 | CH-AWS | SG | 0 | 1 | 0 |  | 0 | 1 | PL: There was a problem on Jan 13. There are a whole load of files with only 29 bytes of data. Seems to have solved it self on Jan 15, and Sonic Show not connected to Irga
18 Jan 2019 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
18 Jan 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
17 Jan 2019 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0 | Eddy working again but with LI7500 (temp. replacement)
15 Jan 2019 | CH-LAE | EPL | 0 | 0 | 2 |  | 2 | 0 | IRGA not working!
14 Jan 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
14 Jan 2019 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | Teamviewer on QCL not running. Frequently there have been problems with QCL teamviewer and QCL internet connection. IT-ticket in progress...
14 Jan 2019 | CH-OE2 | RM | 0 | 0 | 1 |  | 0 | 2 | No complete dataset between 20190108 and 20190114. Restarted irga and should work now.
14 Jan 2019 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
10 Jan 2019 | CH-LAE | TB | 0 | 0 | 1 |  | 0 | 1 | The last few weeks have been an on and off agin from the li7200. Sometimes it worked and sometimes not. We will replace it. See fieldbook
08 Jan 2019 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0
07 Jan 2019 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect directly, but fluxes until end of Dec 2018 look very good and complete which means the logger worked
04 Jan 2019 | CH-INO | MS | 0 | 0 | 0 |  | 0 | 0 | for files 2018122013.XOO - 20182019.XOO data loss due to site maintenance.
04 Jan 2019 | CH-DAV | MG | 0 | 0 | 0 |  | 2 | 2 | No files were generated after 201901031930_L01_F05. I restarted the sonic on 4 January 7:59. Data acquisition in order now.
03 Jan 2019 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0 | All in order
03 Jan 2019 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0 | All in order
20 Dec 2018 | CH-LAE | TB | 0 | 0 | 1 |  | 0 | 0 | The AGC is no very low. Tried to clean the device but didn't help. We need to calibrate the device in 2019 right away.
18 Dec 2018 | CH-AWS | SG | 0 | 0 | 0 |  | 0 | 0 | between 3-5th Dec the file size is not correct
17 Dec 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
17 Dec 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
14 Dec 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
14 Dec 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
14 Dec 2018 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
12 Dec 2018 | CH-AWS | IB | 0 | 0 | 0 |  | 0 | 0
11 Dec 2018 | CH-CHA | IF | 0 | 0 | 1 |  | 0 | 0 | irgashow: says window dirtiness 43.8% TOO LOW. Usually the lowest value is 50%. qclshow: data transfer from QCL now working again after 7 months of absence!!! However qclshow shows the values in the wrong magnitude: QCL CH4 1979690.0 N2O 331059.0 H2O 13273900.0 TEMP 294026.0 PRESS 26510.0 --> divide it by 1000
10 Dec 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
10 Dec 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but I checked all fluxes until 1 Dec and data look OK and complete up until then
07 Dec 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
07 Dec 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
04 Dec 2018 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
03 Dec 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
03 Dec 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
03 Dec 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
03 Dec 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0 | CH-LAS is back up and running. 8 day gap in CH-LAS data.
27 Nov 2018 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
26 Nov 2018 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
26 Nov 2018 | CH-FRU | PL | 0 | 0 | 1 |  | 0 | 0 | Irga H2O values are negative (Window dirtyness 81.2%)
26 Nov 2018 | CH-LAE | EPL | 1 | 1 | 1 |  | 1 | 1 | CH-LAS is offline!!!!
26 Nov 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 2 | 2 | EC files were not generated between 23.11.2018 16:00 and 26.11.2018 09:00 when I restarted the data acquisition.
23 Nov 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
23 Nov 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
23 Nov 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
22 Nov 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but I checked recent fluxes and data look OK and complete up until today
21 Nov 2018 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 1 | Files on 20.11.2018 are smaller because of work in the DAQ (data aquisition). Root partition is nearly full (98%)
21 Nov 2018 | CH-FRU | PL | 0 | 0 | 1 |  | 0 | 0 | AGC is at 93.8%. DAQ root partition is full (100%)
20 Nov 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
19 Nov 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
16 Nov 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
16 Nov 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | lgrshow = 1. Cell pressure is still not constant. At the moment at 28.61 Torr.
16 Nov 2018 | CH-LAE | TB | 0 | 0 | 1 |  | 0 | 0 | AGC is starting to get low. Good we will calibrate soon.
12 Nov 2018 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
12 Nov 2018 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
12 Nov 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0 | No fluxes on quickviews.
09 Nov 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
09 Nov 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
09 Nov 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | lgrshow = 1; cell pressure is still not constantly 45 Torr.
08 Nov 2018 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | Still no fluxes on quickviews
08 Nov 2018 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
07 Nov 2018 | CH-INO | MS | 0 | 0 | 0 |  | 0 | 0 | lgrshow = 1, cell pressure problems, from 30.10 45 torr to 7.11 approx. 40 torr
07 Nov 2018 | CH-OE2 | MS | 0 | 0 | 0 |  | 0 | 0
05 Nov 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
05 Nov 2018 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
05 Nov 2018 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0
05 Nov 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0 | Quickviews for EC fluxes are not working
02 Nov 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 2 | 2 | Except for one file at 01.11.2018 13:30 no EC files were generated between 01.11.2018 9:30 and 02.11.2018 10:30. Files went back to normal after data acquisition was restarted.
02 Nov 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
02 Nov 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
30 Oct 2018 | CH-DAV | MG | 0 | 2 | 0 |  | 2 | 2 | Data gap between 27.10.2018 22:00 and 28.10.2018 10:00. Reason: sonic failure due to heavy snow.
30 Oct 2018 | CH-OE2 | MS | 0 | 0 | 0 |  | 0 | 0
30 Oct 2018 | CH-INO | MS | 0 | 0 | 0 |  | 1 | 0 | In case of pump (external) failure, laser was set to slowflow (1Hz), so filesize decreased accordingly. Affected files 2018102713 to 2018103013
29 Oct 2018 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
29 Oct 2018 | CH-CHA | PL | 0 | 0 | 1 |  | 0 | 0 | AGC low: 43.8%
29 Oct 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
26 Oct 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0 | cell pressure of the LGR is decreasing again
26 Oct 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
26 Oct 2018 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
26 Oct 2018 | CH-DAV | TB | 0 | 0 | 0 |  | 0 | 0
25 Oct 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
24 Oct 2018 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0
24 Oct 2018 | CH-CHA | PL | 0 | 0 | 0 |  | 1 | 1 | File at 13:00 is missing Data. Files on 23.10. at around 9:00 are smaller than expected.
24 Oct 2018 | CH-AWS | IB | 0 | 0 | 0 |  | 0 | 0
23 Oct 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
23 Oct 2018 | CH-INO | MS | 0 | 0 | 0 |  | 0 | 0
23 Oct 2018 | CH-OE2 | MS | 0 | 0 | 0 |  | 0 | 0
22 Oct 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
19 Oct 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
19 Oct 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
19 Oct 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
17 Oct 2018 | CH-DAV | TB | 0 | 0 | 0 |  | 0 | 0
17 Oct 2018 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
16 Oct 2018 | CH-INO | MS | 0 | 0 | 0 |  | 0 | 0
15 Oct 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
15 Oct 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
15 Oct 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
12 Oct 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
12 Oct 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
11 Oct 2018 | CH-DAV | TB | 0 | 0 | 0 |  | 0 | 0
10 Oct 2018 | CH-FRU | PL | 0 | 0 | 0 |  | 0 | 0 | Root partition nearly full (98%). Data partition okay (56%)
10 Oct 2018 | CH-CHA | PL | 0 | 0 | 0 |  | 0 | 0 | No QCL data
09 Oct 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
09 Oct 2018 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
09 Oct 2018 | CH-DAV | TB | 0 | 0 | 0 |  | 0 | 0
09 Oct 2018 | CH-INO | MS | 0 | 0 | 0 |  | 0 | 0
08 Oct 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
08 Oct 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
08 Oct 2018 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | Still gaps in fluxes probably because of dew
05 Oct 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0 | Data gap for almost a week (20180926-20181004, besides 20181002). First, no data was recorded and second, IRGA was calibrated. Since today, everything works again.
05 Oct 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 1 | INO was down for three days due to calibration of the IRGA (20181002-20181004). Still the file problem is not solved yet.
03 Oct 2018 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
01 Oct 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
01 Oct 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
28 Sep 2018 | CH-DAV | MG | 0 | 0 | 1 |  | 0 | 0 | irga flow rate slightly lower than usual (15 instead of 18-19) but still within the normal range.
27 Sep 2018 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0 | LI7200 flow back to normal
26 Sep 2018 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | Dew on Irga, no H2O and CO2 fluxes for time with high window dirtiness
25 Sep 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
25 Sep 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
24 Sep 2018 | CH-LAE | EPL | 0 | 0 | 1 |  | 0 | 0 | Flow of LI-7200 reduced to 8, suggests partial obstruction.
21 Sep 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
21 Sep 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
18 Sep 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 2 | Since Monday I cannot access the FluxData folder (no permission) so I can't see the files for the last week. IT is working on it.
18 Sep 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 2 | Since Monday I cannot access the FluxData folder (no permission) so I can't see the files for the last week. IT is working on it.
17 Sep 2018 | CH-DAV | MG | 0 | 0 | 2 |  | 0 | 0 | No irga data on the Lanner. Instead I get the following message: "Queueid 73. Reading from message queue 32769...(note that this only works under Linux, not MacOS so far)"
16 Sep 2018 | CH-AWS | LH |  |  |  |  | 1 |  | There is a 2-day data gap for EC files, between files 2018091201.A00 and 2018091408.A28. After that looks normal.
16 Sep 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
12 Sep 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
12 Sep 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
10 Sep 2018 | CH-LAE | TB | 0 | 0 | 0 |  | 0 | 0
10 Sep 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
10 Sep 2018 | CH-AWS | TB | 0 | 0 | 0 |  | 0 | 0
05 Sep 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
04 Sep 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
04 Sep 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
04 Sep 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
04 Sep 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
30 Aug 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
28 Aug 2018 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
28 Aug 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
28 Aug 2018 | CH-INO | RM | 0 | 0 | 0 |  | 0 | 0
28 Aug 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
27 Aug 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
22 Aug 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
21 Aug 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
21 Aug 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
14 Aug 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
14 Aug 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
12 Aug 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
10 Aug 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | LH re-started the data storage processes on the Lanner yesterday evening. EC data are being stored OK now.
09 Aug 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
09 Aug 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 2 | 2 | EC data not stored on the Lanner after 201808081400 and therefore not transferred to ETH either. TeamViewer does not work, so I am not able to connect to DAV-MAIN PC.
08 Aug 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
06 Aug 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
02 Aug 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
01 Aug 2018 | CH-OE2 | RM | 1 | 1 | 1 |  | 1 | 0 | Cannot connect to the logger from Germany. Files are arriving on RAW and fluxes on QVs look good.
31 Jul 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
30 Jul 2018 | CH-DAV | LH | 0 | 0 | 0 |  | 0 | 0 | on the LANNER, everything looks good; like yesterday, I was not able to connect to DAV_MAIN PC w/ TeamViewer, but data from the LANNER to ETH was transferred as expected during the night, everything is fine
29 Jul 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
27 Jul 2018 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | Cleaned Sonic transducer and Licor on 27th 09:53-10:00
26 Jul 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
26 Jul 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
24 Jul 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
23 Jul 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
23 Jul 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
23 Jul 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
20 Jul 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
18 Jul 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 1 | 1 | Files stopped growing at 2018071713.X00 and 19:29h (CH-Dav_EC_201807171930_L01_F01.csv) although sonicshow and irgashow were working normally. PM restarted the EC acquisition system and files are gowing again. Additionally, no files in the data or the OK.2018 directory between CH-Dav_EC_201807171500_L01_F01.csv and CH-Dav_EC_201807171930_L01_F01.csv, whereas the file 2018071713.X00 seems to be complete during that time.
17 Jul 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
16 Jul 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
16 Jul 2018 | CH-LAE | EPL | 0 | 0 | 0 |  | 0 | 0
15 Jul 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
13 Jul 2018 | CH-DAV | LH | 0 | 0 | 0 |  | 0 | 0
11 Jul 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
10 Jul 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
06 Jul 2018 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0 | Irga had problems because cable was loose so it only sent data to the logger intermittently. Previous 3-day data loss probably also because of that
05 Jul 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
04 Jul 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
04 Jul 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
02 Jul 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
02 Jul 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | EC files transfer worked again as expected, new EC files are on the grassland server; still cannot connect to the logger (no permission), but will try to find a solution
29 Jun 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0 | Data logging back to normal after an outfall between 21 and 26 June (reason: to be confimed with WE and LH but also data not ingested into archive due to an error in script ALL-Uploads-to-Archive).
26 Jun 2018 | CH-CHA | IF | 0 | 1 | 2 |  | 0 | 0 | No Irga data since June 23. irgashow and sonicshow give message "no IRGA data". --> solved by rebooting the station computer, Irga is running again, 3 days of data loss
26 Jun 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 1 | No EC data between 20180615 and 20180624. But it seems for today they are missing as well.
25 Jun 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 2 | EC data not transfered to ETH after 20180619. Last csv data on LANNER is from 20180621. Storage was full.
19 Jun 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 2 | still cannot connect to Logger, BUT EC fluxes look OK and complete on QVs. However, raw EC files are only available until 14 Jun on GL-RAW, last days are missing
18 Jun 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
12 Jun 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
08 Jun 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
06 Jun 2018 | CH-CHA | IF | 0 | 0 | 0 |  | 0 | 0
05 Jun 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
04 Jun 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
31 May 2018 | CH-OE2 | RM | 0 | 0 | 0 |  | 0 | 0
30 May 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
29 May 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
23 May 2018 | CH-DAV | MG | 0 | 0 | 0 |  | 0 | 0
22 May 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
16 May 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
09 May 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
02 May 2018 | CH-FRU | LH | 1 | 1 | 1 |  | 1 | 1 | still cannot connect to Logger, but EC fluxes look OK and complete on QVs
|  | GA | 0 | 0 | 0 | 68 | 0 | 0
|  | GA | 0 | 0 | 0 | 73 | 0 | 0
|  | GA | 0 | 0 | 0 | 74 | 0 | 0
|  | GA | 0 | 0 | 0 | 74 | 0 | 0
|  | GA | 0 | 0 | 0 | 78 | 0 | 0
|  | GA | 0 | 0 | 0 | 92 | 0 | 0
|  | GA | 0 | 0 | 0 | 48 | 0 | 0
| CH-CHA | YW | 0 | 0 | 0 | 74 | 0 | 0